## 표지(국문)-시흥
ISO/IEC 17025;2017 와 KS Q ISO/IEC 17025;2017 에 의한 공인인정범위
㈜한국교정기술원
경기도 시흥시 서촌상가 4길 17 , 화신오피스텔 제가동 206 (정왕동)
전화 : 031)497-4694 , 팩스 : 031)434-4694 , e-mail : kcti@naver.com
교 정
유효기간 만료일 : 2027. 03. 19 | [6 empty] | 인정번호 : KC07-212호
KOLAS 평가결과에 의해 다음의 교정수행에 대하여 공인됩니다.
분류번호 | 교정항목 | 현장 교정 | 분류번호 | 교정항목 | 현장 교정 | 분류번호 | 교정항목 | 현장 교정
102. 선형치수 |  |  | 10517 | 촉침식 표면 거칠기 측정기 | Y | 20406 | 절대압계; 다이얼, 디지털, | N
10201 | 볼; 강구, 표준구 등 | N | 10525 | 나사 플러그 게이지 | N |  | 기압계, 기록계 등
10206 | 다이얼/실린더 게이지 시험기 | N | 10526 | 테이퍼 나사 플러그 게이지 | N | 20408 | 연성 압력계 | Y
10209 | 엔드바, 마이크로미터 기준봉 | N | 10527 | 나사 링 게이지 | N | 20409 | 차압계; 디지털, 다이얼 포함 | Y
10210 | 길이변위계, LVDT | Y | 10529 | 브이 블록 및 박스 블록 | N | 20411 | 게이지압용 압력계; | Y
10211 | 틈새 게이지 | N | 106. 기타 길이 관련량 |  |  |  | 다이얼, 디지털, 기록계 등
10213 | 갭 게이지 | N | 10601 | 내/외측/기어 이두께 캘리퍼, | Y | 20412 | 압력 변환기/전송기 | Y
10216 | 높이 게이지/측정기 | Y |  | 캘리퍼 게이지 |  | 20413 | 다이얼형 진공계 | Y
10220 | 표준 측장기 | Y | 10603 | 실린더/보어 게이지 | Y | 206. 부피
10223 | 전기 마이크로미터 | N | 10604 | 깊이게이지, 깊이마이크로미터; | Y | 20601 | 유리제 부피계; 타 재질 포함 | N
10224 | 높이 마이크로미터,받침 블록 | N |  | 다이얼형 포함 |  | 20605 | 콘크리트 공기량 시험기 | N
10227 | 표준 내/외경 줄자 | N | 10605 | 다이얼/디지털 게이지 | Y | 20606 | 피스톤식 부피계 | N
10228 | 원통형 플러그/핀 게이지, | N | 10609 | 지침 측미기,테스트 인디케이터 | Y | 207. 밀도
| 나사 측정용 와이어 게이지 |  | 10610 | 마이크로미터 헤드 | N | 20702 | 액체 밀도계 | N
10229 | 래디어스 게이지 | N | 10611 | 3점 마이크로미터 | Y | 20704 | 염도계 | N
10230 | 원통형 링 게이지 | N | 10612 | 내측 마이크로미터 | Y | 20705 | 당도계 | N
10232 | 스텝 게이지/캘리퍼 검사기 | N | 10613 | 외측 마이크로미터 | Y | 20707 | 염화물 측정기 | N
10233 | 테이퍼형 틈새 게이지 | N | 10617 | 표준 체 | N | 208. 점도
10234 | 초음파식 두께 측정기 | Y | 10620 | 용접 게이지 | N | 20801 | 동점도계; 모세관형 점도계 등 | N
10235 | 두께 측정용 기준 시편; 코팅형, | N | 201. 질량 |  |  | 20802 | 회전형 동점도계 | N
| 초음파 시편 |  | 20102 | 자동 호퍼 저울 | Y | 210. 경도
10236 | 피막 두께 측정기 | Y | 20105 | 부등비 접시 수동 저울 | Y | 21001 | 브리넬 경도시험기 | Y
104. 형상 |  |  | 20109 | 전기식 지시 저울 | Y | 21002 | 로크웰 경도시험기 | Y
10401 | 형상 측정기 | Y | 20112 | 판수동/플랫트 폼 저울 | Y | 21003 | 쇼어 경도시험기 | Y
10406 | 평행 블록 | N | 20113 | 접시 지시 저울, 스프링 지시 | Y | 21004 | 비커스 경도시험기 | Y
10407 | 정밀 정반 | Y |  | 저울 |  | 21005 | 듀로미터 경도시험기 | N
10409 | 진원도 측정기 | Y | 20116 | 분동 및 추 | Y | 21006 | 리브 경도시험기 | N
10413 | 곧은자 | N | 202. 힘 |  |  | 301. 시간/주파수
105. 복합형상 |  |  | 20203 | 인장 및 압축 시험기 | Y | 30104 | 주파수 측정기/계수기 | N
10502 | 벤치 센터 | N | 20204 | 푸쉬풀 게이지 | Y | 30106 | 시간간격 측정기, 초시계 및 | Y
10503 | 접촉식 좌표 측정기 | Y | 203. 토크 |  |  |  | 타이머
10504 | 비접촉식 좌표 측정기 | Y | 20302 | 토크 측정기 | Y | 302. 속도/회전수
10511 | 측정현미경, 측정투영기 | Y | 20303 | 토크 렌치 및 토크 드라이버 | Y | 30201 | 표준회전수발생창치 | Y
10512 | 측미 현미경 | N | 204. 압력 |  |  | 30202 | 접촉식 회전 속도계 | N
분류번호 | 교정항목 | 현장 교정 | 분류번호 | 교정항목 | 현장 교정 | 분류번호 | 교정항목 | 현장 교정
30203 | 광 회전 속도계 | N | 501. 접촉식 온도 |  |  | 701. 광도
401. 직류 |  |  | 50101 | 온도 발생장치; 오븐, 전기로, | Y | 70101 | 광 조도계 | N
40101 | 직류 전류계 | Y |  | 액체항온조, 빙점조,
40103 | 직류 전압/전류 교정기 | Y |  | 드라이블럭교정기 등
40104 | 전기식 온도 교정기(센서 미포함) | Y | 50102 | 온도 지시계; 지시/기록/조절계, | Y
40105 | 직류용 분류기 | Y |  | 온도 교정기 등
40108 | 직류 전원 공급기 | Y | 50103 | 유리제 온도계; 유리제 온도계, | N
40112 | 직류 전압계/차동 전압계 등 | Y |  | 벡크만 온도계 등
402.저항, 용량 및 인덕턴스 |  |  | 50104 | 저항식 온도계; | Y
40205 | 접지 저항 측정기 | Y |  | 백금저항온도계, 측온저항체,
40210 | 절연시험기 | Y |  | 써미스터 등
40214 | 저항 측정기, 고저항 측정기 등 | Y | 50105 | 열팽창식 온도계; 바이메탈 | Y
40215 | 저항기, 표준저항, 계단식 저항 | Y |  | 온도계, 기체 또는 액체 충만식
| ,고저항 등 |  |  | 온도계 등
40217 | 임피던스 브리지/LCR 미터 | Y | 50106 | 열전대; 귀금속, 비금속, 순금속, | Y
403.교류 및 교류전력 |  |  |  | 특수 등
40301 | 교류 전류계 | Y | 50107 | 온도 변환기 | Y
40302 | 클램프형 전류계/전압계 | Y | 502. 비접촉식 온도
40303 | 교류 전압/전류 교정기 | Y | 50204 | 복사 온도계 | N
40305 | 교류 전류 분류기 | Y | 503. 습도
40310 | 역률계, 무효율계 등 | Y | 50302 | 상대습도 습도계; | N
40311 | 교류 전력계, 피상.고조파 및 | Y |  | 고분자 박막, 모발 등
| 무효전력계 등 |  | 50303 | 건습구 습도계; 아스만 통풍, | N
40312 | 교류 전원 공급기 | Y |  | 저항온도계식 등
40313 | 내전압/전기 안전 시험기 | Y | 50304 | 온·습도 기록계; | N
40318 | 교류 전압계; 전위차, 실효치 | Y |  | 자기온습도기록계 등
404. 기타 직류 및 저주파 |  |  | 50305 | 노점/상대습도 변환기 | N
40410 | 전원 주파수계 | Y | 50306 | 습도 발생장치; 이압력식/ | Y
40416 | 누설전류 시험기 | Y |  | 이온도식/분류식 습도발생장치,
40417 | 직/교류 전자부하 | Y |  | 항온항습기 등
40419 | 아날로그/디지털 멀티미터; | Y
| 하부 속성 개별 인정 |  | 601. 음향
40421 | 파형측정기 | Y | 60106 | 소음계 | N
주석
1. 위 기관은 고정표준실교정 및 현장교정 서비스를 제공한다.
2. 현장교정 서비스를 제공하는 기관은 "KOLAS-SR-007 현장 시험 및 교정 수행을 위한 추가기술요건"을 만족한다.
3. 인정범위 상에 "Y"가 표기된 항목은 현장교정 서비스가 가능하고, "N"이 표기된 항목은 현장교정 서비스가 불가능하다.
4. 측정불확도는 포함인자 k=2를 사용하며, 신뢰수준 약 95 %를 나타내는 확장불확도로 표현되며 정상적인 조건에서 고객에 제공될 수 있는 최소의 측정불확도를 의미한다.
5. 일반적으로, 교정성적서에 기재된 측정불확도는 교정용 표준기, 고객의 장비와 교정환경 등의 영향 때문에, 공인인정범위에 기재된 측정불확도보다 더 크게 표현됨을 유의하여야 한다.

## 교정측정능력(국문)-시흥
102. 선형치수
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
볼; 강구, 표준구 등 | 10201 | (5 ~ 50) mm | 1.02 + (0.0032 × l02) μm | 표준 측장기
|  |  | (l0:mm) | / KCTI-CQI-10201
다이얼/실린더 게이지 시험기 | 10206 | (0 ~ 100) mm | 0.372 + (0.0022 × l02) μm | 전기마이크로미터
|  |  | (l0:mm) | 게이지블록/ KCTI-CQI-10206
엔드바, 마이크로미터 기준봉 | 10209 | (25 ~ 500) mm | 1.22 + (0.0072 × l02) μm | 전기마이크로미터
|  |  | (l0:mm) | 게이지블록/ KCTI-CQI-10209
길이변위계, LVDT | 10210 | (0 ~ 600) mm | 1.32 + (0.004 22 × l02) μm | 게이지블록,전기 마이크로미터
|  |  | (l0:mm) | / KCTI-CQI-10210
틈새 게이지 | 10211 | (0.01 ~ 3) mm | 1.1 μm | 벤치 마이크로미터
/ KCTI-CQI-10211
갭 게이지 | 10213 | (5 ~ 200) mm | 3.12 + (0.0062 × l02) μm | 전기마이크로미터
|  |  | (l0:mm) | / KCTI-CQI-10213
높이 게이지/측정기 | 10216 | (0 ~ 1 000) mm | 1.62 + (0.0032 × l02) μm | 게이지블록,스텝게이지
|  |  | (l0:mm) | / KCTI-CQI-10216
표준 측장기 | 10220 | (0 ~ 500) mm | 0.32 + (0.0032 × l02) μm | 게이지블록,롱 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10220
전기 마이크로미터 | 10223 | (0 ~ 5) mm | 0.15 μm | 게이지블록
/ KCTI-CQI-10223
높이마이크로미터,받침 블록 | 10224
높이 마이크로미터
블록 |  | (0 ~ 600) mm | 1.32 + (0.005 22 × l02) mm | 전기마이크로미터
|  |  | (l0:m) | 게이지블록/ KCTI-CQI-10224
헤드 |  | (0 ~ 25) mm | 1.42 + (0.0042 × l02) mm
(l0:m)
표준 내/외경 줄자 | 10227
|  | (0 ~ 5) m | 0.072 + (0.0122 × l02) mm | 줄자교정장치
|  |  | (l0:m) | / KCTI-CQI-10227
|  | (5 ~ 50) m | 0.312 + (0.0122 × l02) mm
(l0:m)
원통형 플러그/핀 게이지, 나사 | 10228
측정용 와이어 게이지
원통형 플러그/핀 게이지 |  | (0.1 ~ 100) mm | 2.0 μm | 벤치 마이크로미터
표준 측장기
나사 측정용 와이어 게이지 |  | (0.1 ~ 10) mm | 1.0 μm | / KCTI-CQI-10228
래디어스 게이지 | 10229 | (0 ~ 50) mm | 4.1 μm | 측정 현미경
/ KCTI-CQI-10229
원통형 링 게이지 | 10230 | (3 ~ 150) mm | 1.5 μm | 표준 측장기
/ KCTI-CQI-10230
스텝 게이지/캘리퍼 검사기 | 10232 | (0 ~ 1 000) mm | 1.42 + (0.0042 × l02) μm | 게이지블록,롱 게이지 블록
|  |  | (l0:mm) | 전기마이크로미터
/ KCTI-CQI-10232
테이퍼형 틈새 게이지 | 10233 | (0 ~ 150) mm | 8 μm | 비접촉 좌표측정기
/ KCTI-CQI-10233
102. 선형치수
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
초음파식 두께 측정기 | 10234 | (0 ~ 100) mm | 8 μm | 게이지블록,기준편
/ KCTI-CQI-10234
두께 측정용 기준 시편; 코팅형, | 10235
초음파 시편
코팅형 |  | (1 ~ 5 000) μm | 1.5 μm | 게이지 블록,표준 측장기
/ KCTI-CQI-10235
초음파 시편 |  | (0 ~ 200) mm | 1.02 + (0.0032 × l02) μm
(l0:mm)
피막 두께 측정기 | 10236 | (0 ~ 1.6) mm | 2.5 μm | 표준 두께 시편
/ KCTI-CQI-10236
104. 형상
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
형상 측정기 | 10401
세로방향 (Z축) |  | (0 ~ 100) mm | 0.38 μm | 형상 표준 시편
가로방향 (X축) |  | (1 ~ 50) mm | 0.91 μm | 게이지 블록,각도 게이지블록
각도 |  | (0 ~ 90)˚ | 4.0˝ | / KCTI-CQI-10401
평행 블록 | 10406
평면도 |  | (0 ~ 300) mm | 0.6 μm | 전기 마이크로미터,정밀정반
평행도 |  | (0 ~ 300) mm | 0.6 μm | 높이 게이지
짝(한쌍)의 높이 및 폭의 차이 |  | (0 ~ 1) mm | 0.8 μm | / KCTI-CQI-10406
정밀 정반 | 10407
대각선 길이 |  | (0 ~ 700) mm | 0.7 μm | 수준기
|  | (700 ~ 1 400) mm | 1.4 μm | / KCTI-CQI-10407
|  | (1 400 ~ 2 800) mm | 2.1 μm
|  | (2 800 ~ 4 000) mm | 2.6 μm
|  | (4 000 ~ 6 000) mm | 3.5 μm
진원도 측정기 | 10409
검출기 정확도 |  | (0 ~ 300) μm | 0.40 μm | 진원도 표준 시편
스핀들 원주방향 회전 정확도 |  | (0 ~ 360)˚ | 20 nm | 진원도 표준 반구
스핀들 축방향 회전 정확도 |  | (0 ~ 360)˚ | 40 nm | / KCTI-CQI-10409
곧은자 | 10413
눈금간격 정확도 |  | (0 ~ 3 000) mm | 0.04 mm | 리니어 스케일,틈새게이지
진직도 |  | (0 ~ 3 000) mm | 12 μm | /KCTI-CQI-10413
105. 복합형상
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
벤치 센터 | 10502
양 센터 평행도 |  | (100 ~ 300) mm | 3.4 μm | 게이지 블록
베드면 평면도 |  | (100 ~ 300) mm | 2.2 μm | 전기 마이크로미터
양 센터 높이차 |  | (100 ~ 300) mm | 3.0 μm | / KCTI-CQI-10502
접촉식 좌표 측정기 | 10503
X,Y,Z 지시정확도 |  | (0 ~ 1 000) mm | 0.82 + (0.0042 × l02) μm | 직각자
|  |  | (l0:mm) | 게이지 블록
대각선 |  | (0 ~ 1 000) mm | 0.92 + (0.0042 × l02) μm | 스텝 게이지
|  |  | (l0:mm) | / KCTI-CQI-10503
진직도 |  | (0 ~ 500) mm | 3.0 μm
직각도 |  | (0 ~ 500) mm | 2.0˝
105. 복합형상
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
비접촉식 좌표 측정기 | 10504
지시정확도 |  | (0 ~ 300) mm | 1.1 μm | 표준자,직각자
|  | (300 ~ 500) mm | 1.7 μm | / KCTI-CQI-10504
직각도(Y/X) |  | - | 5˝
측정현미경, 측정투영기 | 10511
측정 현미경 |  | (0 ~ 300) mm | 2.0 μm | 표준자
|  | (300 ~ 500) mm | 3.1 μm | / KCTI-CQI-10511
직각도 |  | (0 ~ 200) mm | 3.1 μm
측정 투영기 |  | (0 ~ 300) mm | 3.5 μm | 표준자
직각도 |  | 100 mm | 2.6 μm | / KCTI-CQI-10511
배율오차 |  | 10 배 | 0.026 %
투영면 회전 각도 |  | (0 ~ 360)˚ | 2´
측미 현미경 | 10512 | (0 ~ 1) mm | 1.0 μm | 표준자
|  | (1 ~ 10) mm | 1.8 μm | / KCTI-CQI-10512
촉침식 표면 거칠기 측정기 | 10517
Ra |  | (0 ~ 3.3) μm | 0.05 μm | 표면 거칠기 시편
Rz |  | (0 ~ 13) μm | 0.18 μm | / KCTI-CQI-10517
RSm |  | (0 ~ 1) mm | 0.06 μm
단차높이 |  | (0 ~ 10) μm | 0.20 μm
나사 플러그 게이지 | 10525
유효 지름 |  | (1 ~ 50) mm | 1.9 μm | 벤치 마이크로미터
피치 |  | (0 ~ 3) mm | 1.3 μm | 나사측정용 와이어 게이지
바깥 지름 |  | (0 ~ 50) mm | 1.5 μm | 비접촉식 좌표 측정기
산의 반각 |  | (0 ~ 60)˚ | 2.0´ | / KCTI-CQI-10525
테이퍼 나사 플러그 게이지 | 10526
소단부 유효지름 |  | (4 ~ 50) mm | 5.7 μm | 게이지 블록,표준 측장기
대단부 유효지름 |  | (4 ~ 50) mm | 5.7 μm | 전기마이크로미터,
소단부 바깥지름 |  | (4 ~ 50) mm | 5.7 μm | 나사측정용 와이어게이지
대단부 바깥지름 |  | (4 ~ 50) mm | 5.7 μm | 높이 마이크로미터
피치 |  | (0.5 ~ 4) mm | 2.1 μm | 비접촉식 좌표 측정기
산의 반각 |  | (0 ~ 45)˚ | 1´ 40˝ | / KCTI-CQI-10526
테이퍼의 반각 |  | (0 ~ 45)˚ | 58˝
게이지의 길이 |  | (0 ~ 100) mm | 2.0 μm
나사 링 게이지 | 10527
유효 지름 |  | (4 ~ 50) mm | 2.3 μm | 표준측장기,3점 마이크로미터
안지름 |  | (4 ~ 50) mm | 2.2 μm | / KCTI-CQI-10527
브이 블록 및 박스 블록 | 10529
밑면의 평면도 |  | (0 ~ 150) mm | 1.8 μm | 전기 마이크로미터
브이면의 평면도 |  | (0 ~ 150) mm | 1.8 μm | 테스트 바
밑면과 브이면 위의 |  | (0 ~ 150) mm | 3.8 μm | / KCTI-CQI-10529
원통과의 평행도
브이 홈의 밑면에 대한 기울기 |  | (0 ~ 150) mm | 1.3 μm
옆면과 브이면 위측 |  | (0 ~ 150) mm | 3.8 μm
원통과의 평행도
1쌍의 브이 블록에 대한 |  | (0 ~ 150) mm | 3.8 μm
브이면의 높이의 상호차
106. 기타 길이 및 관련량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
내/외측/기어 이두께 캘리퍼, | 10601
캘리퍼 게이지
내/외측 캘리퍼 |  | (0 ~ 1 500) mm | 102 + (0.0072 × l02) μm | 캘리퍼검사기,게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10601
캘리퍼 게이지 |  | (0 ~ 100) mm | 1.12 + (0.0042 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10601
실린더/보어 게이지 | 10603 | (0 ~ 1.2) mm | 0.7 μm | 다이얼게이지시험기
/ KCTI-CQI-10603
깊이게이지, 깊이마이크로미터; | 10604
다이얼형 포함
깊이 게이지,다이얼형 포함 |  | (0 ~ 500) mm | 82 + (0.0042 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10604
깊이 마이크로미터 |  | (0 ~ 150) mm | 2.02 + (0.0072 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10604
다이얼/디지털 게이지 | 10605 | (0 ~ 25) mm | 1.0 μm | 다이얼게이지시험기
|  | (25 ~ 100) mm | 1.6 μm | / KCTI-CQI-10605
지침 측미기, 테스트 인디케이터 | 10609
지침 측미기 |  | (0 ~ 1) mm | 0.7 μm | 다이얼게이지시험기
/ KCTI-CQI-10609
테스트 인디케이터 |  | (0 ~ 1) mm | 0.7 μm | 다이얼게이지시험기
/ KCTI-CQI-10609
마이크로미터 헤드 | 10610 | (0 ~ 100) mm | 0.82 + (0.0072 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10610
3점 마이크로미터 | 10611 | (3 ~ 5) mm | 1.6 μm | 원통형 링 게이지
|  | (5 ~ 100) mm | 2.3 μm | / KCTI-CQI-10611
|  | (100 ~ 150) mm | 2.5 μm
내측 마이크로미터 | 10612
내측 마이크로미터 |  | (5 ~ 200) mm | 1.22 + (0.0072 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10612
봉형 마이크로미터 |  | (50 ~ 500) mm | 1.32 + (0.0042 × l02) μm | 게이지블록,표준 측장기
|  |  | (l0:mm) | / KCTI-CQI-10612
외측 마이크로미터 | 10613
외측 마이크로미터 |  | (0 ~ 25) mm | 0.52 + (0.0022 × l02) μm | 게이지블록
|  |  | (l0:mm) | 롱 게이지블록
|  | (25 ~ 125) mm | 1.02 + (0.0032 × l02) μm | / KCTI-CQI-10613
(l0:mm)
|  | (125 ~ 500) mm | 1.22 + (0.0052 × l02) μm
(l0:mm)
V-앤빌 마이크로미터 |  | (2.5 ~ 25) mm | 1.3 μm | 핀 게이지
/ KCTI-CQI-10613
표준체 | 10617
선재의 지름 |  | (0.02 ~ 10) mm | 3.2 μm | 비접촉식 좌표 측정기
체눈의 크기 |  | (0.02 ~ 100) mm | 4.8 μm | / KCTI-CQI-10617
106. 기타 길이 및 관련량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
용접 게이지 | 10620
길이 |  | (0 ~ 150) mm | 0.2 mm | 비접촉식 좌표 측정기
/ KCTI-CQI-10620
각도 |  | (0 ~ 90)˚ | 0.2˚
201. 질량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
자동 호퍼 저울 | 20102
|  | (0 ~ 10) kg | 1.5 g | 분동
|  | (10 ~ 50) kg | 8.1 g | / KCTI-CQI-20102
|  | (50 ~ 200) kg | 74 g
|  | (200 ~ 500) kg | 0.31 kg
부등비 접시 수동 저울 | 20105 | (0 ~ 310) g | 29 mg | 분동
|  | (310 ~ 2 610) g | 62 mg | / KCTI-CQI-20105
|  | (2 610 ~ 20 000) g | 8.6 g
전기식 지시 저울 | 20109 | (0 ~ 200) g | 0.20 mg | 분동
|  | (200 ~ 500) g | 0.51 mg | / KCTI-CQI-20109
|  | (500 ~ 1 000) g | 0.53 mg
|  | (1 ~ 3) kg | 2.0 mg
|  | (3 ~ 5) kg | 2.8 mg
|  | (5 ~ 10) kg | 10 mg
|  | (10 ~ 30) kg | 77 mg
|  | (30 ~ 50) kg | 0.75 g
|  | (50 ~ 200) kg | 1.5 g
|  | (200 ~ 300) kg | 4.4 g
|  | (300 ~ 500) kg | 7.0 g
|  | (500 ~ 1 000) kg | 45 g
|  | (1 000 ~ 3 000) kg | 85 g
판 수동/플랫트 폼 저울 | 20112 | (0 ~ 50) kg | 9.1 g | 분동
|  | (50 ~ 200) kg | 91 g | / KCTI-CQI-20112
|  | (200 ~ 500) kg | 0.19 kg
접시 지시 저울, 스프링 지시저울 | 20113
|  | (0 ~ 1) kg | 0.64 g | 분동
|  | (1 ~ 5) kg | 3.1 g | / KCTI-CQI-20113
|  | (5 ~ 10) kg | 6.2 g
|  | (10 ~ 50) kg | 58 g
|  | (50 ~ 100) kg | 0.31 kg
201. 질량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
분동 및 추 | 20116 | (1 mg ~ 10 kg) | F1급 이하 | 분동,전기식 지시저울
|  | 1 mg | 2.6 μg | / KCTI-CQI-20116
|  | 2 mg | 2.4 μg
|  | 5 mg | 2.4 μg
|  | 10 mg | 3.0 μg
|  | 20 mg | 3.5 μg
|  | 50 mg | 4.3 μg
|  | 100 mg | 5.3 μg
|  | 200 mg | 6.2 μg
|  | 500 mg | 8.2 μg
|  | 1 g | 9.1 μg
|  | 2 g | 12 μg
|  | 5 g | 15 μg
|  | 10 g | 18 μg
|  | 20 g | 24 μg
|  | 50 g | 0.21 mg
|  | 100 g | 0.21 mg
|  | 200 g | 0.16 mg
|  | 500 g | 0.30 mg
|  | 1 kg | 0.6 mg
|  | 2 kg | 2.3 mg
|  | 5 kg | 3.3 mg
|  | 10 kg | 6.4 mg
|  | (20 kg) | M1급 이하
|  | 20 kg | 0.15 g
202. 힘
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
인장 및 압축 시험기 | 20203
인장 |  | (20 ~ 200) N | 1.6 × 10-3 | 분동,전기식 힘 측정기
|  | (0.2 ~ 5) kN | 1.2 × 10-3 | / KCTI-CQI-20203
|  | (5 ~ 10) kN | 1.2 × 10-3
압축 |  | (20 ~ 500) N | 1.3 × 10-3 | 분동,전기식 힘 측정기
|  | (0.5 ~ 20) kN | 1.3 × 10-3 | / KCTI-CQI-20203
|  | (20 ~ 50) kN | 1.8 × 10-3
|  | (50 ~ 100) kN | 2.0 × 10-3
|  | (100 ~ 200) kN | 2.3 × 10-3
|  | (200 ~ 500) kN | 1.7 × 10-3
|  | (0.5 ~ 1) MN | 2.4 × 10-3
|  | (1 ~ 2) MN | 3.2 × 10-3
|  | (2 ~ 3) MN | 1.7 × 10-3
푸쉬풀 게이지 | 20204 | (0.4 ~ 500) N | 1.1 × 10-3 | 추
|  | (500 ~ 1 000) N | 7.5 × 10-4 | / KCTI-CQI-20204
203. 토크
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
토크 측정기 | 20302 | (0.1 ~ 1) N·m | 2.8 × 10-3 | 토크암,추
|  | (1 ~ 10) N·m | 2.6 × 10-3 | / KCTI-CQI-20302
|  | (10 ~ 100) N·m | 1.8 × 10-3
|  | (100 ~ 200) N·m | 1.4 × 10-3
|  | (200 ~ 500) N·m | 1.3 × 10-3
|  | (500 ~ 1 000) N·m | 1.7 × 10-3
|  | (1 000 ~ 1 500) N·m | 1.7 × 10-3
토크 렌치 및 토크 드라이버 | 20303 | (0.2 ~ 50) N·m | 1.0 × 10-2 | 토크측정기
|  | (50 ~ 200) N·m | 1.1 × 10-2 | / KCTI-CQI-20303
|  | (200 ~ 500) N·m | 1.0 × 10-2
204. 압력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
절대압계; 다이얼,디지털, | 20406
기압계, 기록계 등
기압계 |  | (80 ~ 120) kPa | 2.8 × 10 -4 | 디지털 압력계
/ KCTI-CQI-20406
연성압력계 | 20408 | (-95 ~ 2 500) kPa | 7.5 × 10 -4 | 디지털 압력계
/ KCTI-CQI-20408
차압계; 디지털, 다이얼 포함 | 20409 | (0 ~ 0.2) kPa | 5.0 × 10 -3 | 디지털 압력계
|  | (0.2 ~ 2.5) kPa | 7.0 × 10 -4 | 공압 분동식 압력계
|  | (2.5 ~ 600) kPa | 4.0 × 10 -3 | / KCTI-CQI-20409
게이지압용 압력계; 다이얼, | 20411
디지털,기록계 등
다이얼, 디지털 압력계, |  | (0 ~ 0.1) MPa | 2.0 × 10 -4 | 공압/유압 분동식 압력계
기록계 등 |  | (0.1 ~ 100) MPa | 1.0 × 10 -4 | 디지털 압력계
/ KCTI-CQI-20411
압력변환기/전송기 | 20412 | (0 ~ 0.1) MPa | 7.4 × 10 -4 | 공압/유압 분동식 압력계
|  | (0.1 ~ 100) MPa | 7.2 × 10 -4 | 멀티미터,디지털 압력계
/ KCTI-CQI-20412
다이얼형 진공계 | 20413 | (-95 ~ 0) kPa | 2.2 × 10 -3 | 디지털 압력계
/ KCTI-CQI-20413
206. 부피
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
유리제 부피계; 타 재질 포함 | 20601 | (0 ~ 1) ml | 0.63 μl | 전기식 지시저울
|  | (1 ~ 2) ml | 0.95 μl | / KCTI-CQI-20601
|  | (2 ~ 5) ml | 1.7 μl
|  | (5 ~ 10) ml | 2.3 μl
|  | (10 ~ 25) ml | 4.2 μl
|  | (25 ~ 50) ml | 7.0 μl
|  | (50 ~ 100) ml | 14 μl
|  | (100 ~ 200) ml | 29 μl
|  | (200 ~ 250) ml | 55 μl
|  | (250 ~ 500) ml | 86 μl
|  | (500 ~ 1 000) ml | 0.17 ml
|  | (1 000 ~ 2 000) ml | 0.33 ml
콘크리트 공기량 시험기 | 20605 | (0 ~ 10 ) % | 0.05 % | 전기식 지시저울,분동
/ KCTI-CQI-20605
피스톤식 부피계 | 20606 | (0.001 ~ 0.01) ml | 18 nl | 전기식 지시저울
|  | (0.01 ~ 0.02) ml | 40 nl | / KCTI-CQI-20606
|  | (0.02 ~ 0.05) ml | 70 nl
|  | (0.05 ~ 0.1) ml | 92 nl
|  | (0.1 ~ 0.2) ml | 0.20 μl
|  | (0.2 ~ 0.5) ml | 0.36 μl
|  | (0.5 ~ 1) ml | 0.66 μl
|  | (1 ~ 2) ml | 1.3 μl
|  | (2 ~ 5) ml | 3.1 μl
|  | (5 ~ 10) ml | 6.1 μl
|  | (10 ~ 20) ml | 13 μl
|  | (20 ~ 50) ml | 31 μl
207. 밀도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
액체 밀도계 | 20702 | (0.75 ~ 1.5) g/cm3 | 0.000 10 g/cm3 | 밀도 표준물,비중병
/ KCTI-CQI-20702
염도계 | 20704 | (0 ~ 26) % | 0.042% | NaCl
/ KCTI-CQI-20704
당도계 | 20705 | (0 ~ 60) % | 0.044% | Sucrose
/ KCTI-CQI-20705
염화물 측정기 | 20707 | (0 ~ 0.5) % | 0.0008% | 밀도 표준물
|  | (0.5 ~ 1.0) % | 0.0009% | / KCTI-CQI-20707
|  | (1.0 ~ 1.5) % | 0.0011%
208. 점도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
동점도계; 모세관형 점도계 등 | 20801
잔컵 점도계 |  | (2.5 ~ 1 000) mm2/s | 2.6 × 10-2 | 점도액
/ KCTI-CQI-20801
포드컵 점도계 |  | (2.5 ~ 1 000) mm2/s | 2.6 × 10-2
회전형 동점도계 | 20802
회전형 점도계 |  | (20 ~ 200 000) mPa·s | 1.6 × 10-2 | 점도액
/ KCTI-CQI-20802
210. 경도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
브리넬 경도시험기 | 21001 | (100 ~ 250) HBW 10/3 000 | 3.0 HBW 10 / 3 000 | 브리넬경도기준편
|  | (250 ~ 450) HBW 10/3 000 | 4.9 HBW 10 / 3 000 | 측정현미경,전기식힘측정기
표준자 / KCTI-CQI-21001
로크웰 경도시험기 | 21002 | (20 ~ 70) HRC | 0.46 HRC | 로크웰경도기준편
|  | (20 ~ 100) HRBW | 0.72 HRBW | 측정현미경,전기식힘측정기
인덱스마스터
/ KCTI-CQI-21002
쇼어 경도 시험기 | 21003 | (20 ~ 100) HS | 1.1 HS | 쇼어경도기준편
/ KCTI-CQI-21003
비커스 경도시험기 | 21004 | (40 ~ 240) HV 0.2 | 8.2 HV 0.2 | 비커스경도기준편
|  | (240 ~ 600) HV 0.2 | 18 HV 0.2 | 측정현미경,전기식힘측정기
|  | (600 ~ 750) HV 0.2 | 39 HV 0.2 | 표준자
/ KCTI-CQI-21004
듀로미터 경도시험기 | 21005 | (0 ~ 100) HDA | 1.4 HDA | 듀로미터 시험기
|  | (0 ~ 100) HDD | 1.1 HDD | / KCTI-CQI-21005
리브 경도시험기 | 21006 | 500 HLD 미만 | 4.6 | 리브경도기준편
|  | (500 ~ 700) HLD | 4.6 | / KCTI-CQI-21006
|  | 700 HLD 초과 | 4.6
301. 시간/주파수
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
주파수 측정기/계수기 | 30104
표준주파수와 비교 |  | 10 MHz | 7.2 × 10 -12 | GPS수신기,주파수 계수기
/ KCTI-CQI-30104
기준 발진기 |  | 10 MHz | 8.0 × 10 -13
시간간격 측정기,초시계 및 | 30106
타이머
초시계 및 타이머 |  | 100 ms ~ 86 400 s | 4.2 × 10 -7 | 초시계 교정기
/ KCTI-CQI-30106
302. 속도/회전수
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
표준 회전수 발생장치 | 30201
회전속도 |  | (6 ~ 60) min-1 | 1.2 min-1 | GPS수신기, 주파수계수기,
|  | (60 ~ 600) min-1 | 0.13 min-1 | 표준 타코메타
|  | (600 ~ 6 000) min-1 | 0.44 min-1 | / KCTI-CQI-30201
|  | (6 000 ~ 60 000) min-1 | 4.3 min-1
|  | (60 000 ~ 90 000) min-1 | 6.4 min-1
접촉식 회전 속도계 | 30202 | (6 ~ 4 000) min-1 | 0.083 min-1 | GPS수신기,RPM교정 시스템
/ KCTI-CQI-30202
광 회전 속도계 | 30203
광 회전 속도계 |  | (6 ~ 6 000) min-1 | 0.064 min-1 | GPS수신기,광회전 신호검출기
|  | (6 000 ~ 90 000) min-1 | 0.64 min-1 | / KCTI-CQI-30203
섬광 발생기 |  | (6 ~ 6 000) min-1 | 0.022 min-1
|  | (6 000 ~ 90 000) min-1 | 0.63 min-1
401. 직류
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
직류 전류계 | 40101
직류 전류 |  | 1 mA | 1.7 × 10 -4 | Meter Calibrator
|  | (1 ~ 10) mA | 1.5 × 10 -4 | / KCTI-CQI-40101
|  | (10 ~ 100) mA | 1.5 × 10 -4
|  | (0.1 ~ 1) A | 2.8 × 10 -4
|  | (1 ~ 10) A | 6.4 × 10 -4
직류 전압/전류 교정기 | 40103
직류전압 |  | (0 ~ 100) mV | 1.1 × 10 -5 | Digital Multimeter
|  | (0.1 ~ 1) V | 6.3 × 10 -6 | Current Shunt
|  | (1 ~ 10) V | 5.3 × 10 -6 | / KCTI-CQI-40103
|  | (10 ~ 100) V | 7.7 × 10 -6
|  | (100 ~ 1 000) V | 8.0 × 10 -6
직류전류 |  | (0 ~ 100) μA | 9.6 × 10 -5
|  | (0.1 ~ 1) mA | 3.7 × 10 -5
|  | (1 ~ 10) mA | 3.5 × 10 -5
|  | (10 ~ 100) mA | 6.3 × 10 -5
|  | (0.1 ~ 1) A | 2.3 × 10 -4
|  | (1 ~ 10) A | 5.2 × 10 -4
|  | (10 ~ 20) A | 4.9 × 10 -4
|  | (20 ~ 100) A | 2.0 × 10 -4
전기식 온도 교정기(센서 미포함) | 40104
열전대(SOURCE) |  | (-8.825 ~ 76.373) mV | 1.1 μV | Meter Calibrator
Digital Multimeter
측온저항체 (SOURCE) |  | (17.136 ~ 100)Ω | 0.92 mΩ | / KCTI-CQI-40104
|  | (0.1 ~ 0.5)kΩ | 5.0 mΩ
|  | (0.5 ~ 1)kΩ | 9.2 mΩ
|  | (1 ~ 3.137 1)kΩ | 34 mΩ
열전대(MEASURE) |  | (-8.825 ~ 76.373) mV | 3.8 μV
측온저항체 (MEASURE) |  | (17.136 ~ 100)Ω | 13 mΩ
|  | (0.1 ~ 0.5)kΩ | 70 mΩ
|  | (0.5 ~ 1)kΩ | 82 mΩ
|  | (1 ~ 3.137 1)kΩ | 0.66 Ω
401. 직류
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
직류용 분류기 | 40105
저항 |  | (0 ~ 1) mΩ | 1.5 × 10 -4 | Digital Multimeter
|  | (1 ~ 10) mΩ | 1.3 × 10 -3 | Meter Calibrator
|  | (10 ~ 100) mΩ | 1.2 × 10 -3 | Current Amplifier
|  | (0.1 ~ 1) Ω | 1.5 × 10 -4 | / KCTI-CQI-40105
|  | (1 ~ 10) Ω | 1.5 × 10 -4
|  | (10 ~ 100) Ω | 1.8 × 10 -4
|  | (0.1 ~ 1) kΩ | 7.0 × 10 -4
직류 전원 공급기 | 40108
직류전압 |  | (0 ~ 100) mV | 5.9 × 10 -5 | Meter Calibrator
|  | (0.1 ~ 1) V | 5.8 × 10 -5 | / KCTI-CQI-40108
|  | (1 ~ 10) V | 7.8 × 10 -6
|  | (10 ~ 100) V | 9.6 × 10 -6
|  | (100 ~ 1 000) V | 5.9 × 10 -5
직류전류 |  | (0 ~ 100) μA | 5.9 × 10 -4
|  | (0.1 ~ 1) mA | 5.8 × 10 -3
|  | (1 ~ 10) mA | 5.8 × 10 -4
|  | (10 ~ 100) mA | 8.6 × 10 -5
|  | (0.1 ~ 1) A | 1.2 × 10 -3
|  | (1 ~ 10) A | 1.5 × 10 -4
|  | (10 ~ 100) A | 2.0 × 10 -4
직류 전압계/차동 전압계 등 | 40112
직류 전압 |  | (1 ~ 100) mV | 3.8 × 10 -5 | Meter Calibrator
|  | (0.1 ~ 1) V | 1.8 × 10 -5 | / KCTI-CQI-40112
|  | (1 ~ 10) V | 1.8 × 10 -5
|  | (10 ~ 100) V | 2.4 × 10 -5
|  | (100 ~ 1 000) V | 2.6 × 10 -5
402. 저항, 용량 및 인덕턴스
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
접지 저항 측정기 | 40205
직류저항 |  | (0 ~ 100) mΩ | 4.1 × 10 -3 | High Voltage Meter
|  | (0.1 ~ 1) Ω | 7.1 × 10 -4 | Decade Resistance Box
|  | (0.001 ~ 100) kΩ | 7.1 × 10 -4 | Meter Calibrator
/ KCTI-CQI-40205
교류전압 |  | 50 Hz ~ 100 Hz
|  | (1 ~ 100) V | 5.8 × 10 -3
|  | (100 ~ 300) V | 3.7 × 10 -4
절연시험기 | 40210
절연저항 |  | (1 ~ 10) kΩ | 6.2 × 10 -4 | High Voltage Meter
|  | (10 ~ 100) kΩ | 6.2 × 10 -4 | Decade Resistance Box
|  | (0.1 ~ 1) MΩ | 6.2 × 10 -4 | Meter Calibrator
|  | (1 ~ 10) MΩ | 7.2 × 10 -4 | / KCTI-CQI-40210
|  | (10 ~ 100) MΩ | 7.2 × 10 -4
|  | (0.1 ~ 1) GΩ | 8.2 × 10 -3
|  | (1 ~ 10) GΩ | 1.3 × 10 -2
|  | (10 ~ 100) GΩ | 1.3 × 10 -2
|  | (0.1 ~ 1) TΩ | 1.3 × 10 -2
교류전압 |  | 60 Hz
|  | (10 ~ 100) V | 5.8 × 10 -3
|  | (100 ~ 200) V | 6.2 × 10 -4
|  | (200 ~ 300) V | 3.7 × 10 -4
|  | (300 ~ 400) V | 4.3 × 10 -4
|  | (400 ~ 500) V | 4.3 × 10 -4
|  | (500 ~ 600) V | 4.0 × 10 -4
|  | (600 ~ 1 000) V | 6.8 × 10 -4
402. 저항, 용량 및 인덕턴스
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
절연시험기 | 40210
직류출력전압 |  | (1 ~ 10) V | 6.0 × 10 -5 | High Voltage Meter
|  | (10 ~ 50) V | 6.0 × 10 -5 | Decade Resistance Box
|  | (50 ~ 100) V | 6.0 × 10 -5 | Meter Calibrator
|  | (100 ~ 500) V | 6.0 × 10 -5 | / KCTI-CQI-40210
|  | (0.5 ~ 1) kV | 6.2 × 10 -5
|  | (1 ~ 5) kV | 6.2 × 10 -3
저항측정기, 고저항측정기등 | 40214
저항 |  | 1 mΩ | 6.4 × 10 -5 | Decade Resistance Box
|  | (1 ~ 10) mΩ | 6.4 × 10 -5 | Standard Resistor
|  | (10 ~ 100) mΩ | 2.1 × 10 -5 | / KCTI-CQI-40214
|  | (0.1 ~ 1) Ω | 1.6 × 10 -5
|  | (1 ~ 10) Ω | 1.3 × 10 -5
|  | (10 ~ 100) Ω | 1.2 × 10 -5
|  | (0.1 ~ 1) kΩ | 1.1 × 10 -5
|  | (1 ~ 10) kΩ | 1.1 × 10 -5
|  | (10 ~ 100) kΩ | 1.4 × 10 -4
|  | (0.1 ~ 1) MΩ | 6.2 × 10 -4
|  | (1 ~ 10) MΩ | 7.2 × 10 -4
|  | (10 ~ 100) MΩ | 7.2 × 10 -4
저항기,표준저항,계단식저항 | 40215
고저항 등
저항 |  | (0 ~ 10) mΩ | 1.1 × 10 -3 | Digital Multimeter
|  | (10 ~ 100) mΩ | 5.0 × 10 -4 | / KCTI-CQI-40215
|  | (0.1 ~ 1) Ω | 1.6 × 10 -4
|  | (1 ~ 10) Ω | 1.8 × 10 -5
|  | (10 ~ 100) Ω | 1.1 × 10 -5
|  | (0.1 ~ 100) kΩ | 1.1 × 10 -5
|  | (0.1 ~ 1) MΩ | 1.3 × 10 -5
|  | (1 ~ 10) MΩ | 5.0 × 10 -5
|  | (10 ~ 100) MΩ | 4.3 × 10 -4
|  | (0.1 ~ 1) GΩ | 4.7 × 10 -3
임피던스 브리지/LCR 미터 | 40217
Resistance |  | (0 ~ 1) Ω | 7.3 × 10 -4 | Standard Capacitor
(1 kHz) |  | (1 ~ 10) Ω | 6.8 × 10 -4 | Standard Inductor
|  | (10 ~ 100) Ω | 6.5 × 10 -4 | / KCTI-CQI-40217
|  | (0.1 ~ 1) kΩ | 2.9 × 10 -4
|  | (1 ~ 100) kΩ | 1.2 × 10 -4
|  | (0.1 ~ 1) MΩ | 2.9 × 10 -4
Capacitance |  | (0 ~ 1 000) pF | 3.5 × 10 -4
(1 kHz) |  | (1 ~ 10) nF | 7.6 × 10 -5
|  | (0.01 ~ 1) μF | 1.1 × 10 -4
Inductance |  | (0 ~ 100) μH | 1.2 × 10 -3
(1 kHz) |  | 100 μH ~ 1 H | 3.1 × 10 -4
403. 교류 및 교류전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
교류 전류계 | 40301
교류전류 |  | 50 Hz ~ 1 kHz
|  | (1 ~ 100) mA | 1.4 × 10 -3 | Meter Calibrator
|  | (0.1 ~ 1) A | 8.5 × 10 -4 | / KCTI-CQI-40301
|  | (1 ~ 10) A | 2.4 × 10 -3
클램프형 전류계/전압계 | 40302
직류전압 |  | (10 ~ 100) mV | 1.6 × 10 -4 | Current Coil
|  | (0.1 ~ 1) V | 3.8 × 10 -5 | Meter Calibrator
|  | (1 ~ 1 000) V | 2.4 × 10 -5 | / KCTI-CQI-40302
교류전압 |  | 40 Hz
|  | (10 ~ 100) mV | 1.8 × 10 -3
|  | (0.1 ~ 1) V | 4.5 × 10 -4
|  | (1 ~ 10) V | 4.3 × 10 -4
50 Hz ~ 1 kHz
|  | (10 ~ 100) mV | 1.1 × 10 -3
|  | (0.1 ~ 1) V | 2.8 × 10 -4
|  | (1 ~ 1 000) V | 3.7 × 10 -4
직류전류 |  | (1 ~ 100) mA | 1.9 × 10 -4
|  | (0.1 ~ 10) A | 1.3 × 10 -3
|  | (10 ~ 1 000) A | 3.0 × 10 -3
교류전류 |  | 50 Hz ~ 1 kHz
|  | (1 ~ 100) mA | 1.4 × 10 -3
|  | (0.1 ~ 1) A | 8.5 × 10 -4
|  | (1 ~ 10) A | 2.4 × 10 -3
50 Hz ~ 100 Hz
|  | (1 ~ 100) mA | 1.4 × 10 -3
|  | (0.1 ~ 1) A | 8.5 × 10 -4
|  | (1 ~ 10) A | 2.0 × 10 -3
|  | (10 ~ 1 000) A | 5.4 × 10 -3
저항 |  | (1 ~ 100) Ω | 1.2 × 10 -3
|  | (0.1 ~ 100) kΩ | 1.3 × 10 -4
|  | (0.1 ~ 10) MΩ | 1.8 × 10 -4
|  | (10 ~ 100) MΩ | 6.4 × 10 -4
교류 전압/전류 교정기 | 40303
교류전압 |  | 40 Hz ~ 100 Hz
|  | (10 ~ 100) mV | 2.6 × 10 -4 | Digital Multimeter
|  | (0.1 ~ 1) V | 1.6 × 10 -4 | Current Shunt
|  | (1 ~ 10) V | 1.6 × 10 -4 | Frequency Counter
|  | (10 ~ 100) V | 1.6 × 10 -4 | / KCTI-CQI-40303
50 Hz ~ 100 Hz
|  | (100 ~ 1 000) V | 1.5 × 10 -4
403. 교류 및 교류전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
교류 전압/전류 교정기 | 40303 | 100 Hz ~ 1 kHz |  | Digital Multimeter
|  | (10 ~ 100) mV | 2.2 × 10 -4 | Current Shunt
|  | (0.1 ~ 1) V | 1.2 × 10 -4 | Frequency Counter
|  | (1 ~ 10) V | 1.2 × 10 -4 | / KCTI-CQI-40303
|  | (10 ~ 100) V | 1.2 × 10 -4
|  | (100 ~ 1 000) V | 1.5 × 10 -4
교류전류 |  | 40 Hz ~ 100 Hz
|  | (0.1 ~ 1) mA | 5.6 × 10 -4
|  | (1 ~ 100) mA | 5.6 × 10 -4
|  | (0.1 ~ 1) A | 1.1× 10 -3
|  | (1 ~ 10) A | 1.3× 10 -3
|  | (10 ~ 100) A | 3.3× 10 -3
50 Hz ~ 100 Hz
|  | (10 ~ 100) A | 3.3× 10 -3
100 Hz ~ 1 kHz
|  | (0.1 ~ 1) mA | 5.6 × 10 -4
|  | (1 ~ 100) mA | 5.6 × 10 -4
|  | (0.1 ~ 1) A | 1.1× 10 -3
|  | (1 ~ 10) A | 1.4× 10 -3
|  | (10 ~ 100) A | 3.3× 10 -3
교류 전류 분류기 | 40305
교류저항 |  | 60 Hz ~ 1 kHz |  | Digital Multimeter
|  | (1 ~ 10) mΩ | 2.4 × 10 -3 | Current Shunt
|  | (10 ~ 100) mΩ | 8.8 × 10 -4 | Frequency Counter
|  | (0.1 ~ 1) Ω | 7.6 × 10 -4 | / KCTI-CQI-40305
|  | (1 ~ 10) Ω | 7.6 × 10 -4
|  | (10 ~ 100) Ω | 1.4 × 10 -3
|  | (0.1 ~ 1) kΩ | 3.0 × 10 -3
역률계, 무효율계 등 | 40310
역률 |  | 50 Hz, 60 Hz |  | Meter Calibrator
(LEAD & LAG) |  | 1 | 0.001 | / KCTI-CQI-40310
|  | 0.9 | 0.002
|  | 0.8 | 0.002
|  | 0.7 | 0.003
|  | 0.6 | 0.003
|  | 0.5 | 0.004
|  | 0.4 | 0.005
|  | 0.3 | 0.007
|  | 0.2 | 0.01
|  | 0.1 | 0.021
교류 전력계, 피상.고조파 및 | 40311
무효전력계 등
교류전압 |  | 50 Hz, 60 Hz |  | Meter Calibrator
|  | 10 mV | 1.0 × 10 -3 | / KCTI-CQI-40311
|  | (10 ~ 100) mV | 2.7 × 10 -4
|  | (0.1 ~ 1) V | 2.5 × 10 -4
|  | (1 ~ 10) V | 2.5 × 10 -4
|  | (10 ~ 100) V | 2.5 × 10 -4
|  | (100 ~ 1 000) V | 3.6 × 10 -4
교류전류 |  | 50 Hz, 60 Hz
|  | 100 μA | 2.7 × 10 -3
|  | (0.1 ~ 1) mA | 1.4 × 10 -3
403. 교류 및 교류전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
교류 전력계, 피상.고조파 및 | 40311
무효전력계 등
교류전류 |  | (1 ~ 10) mA | 7.0 × 10 -4 | Meter Calibrator
|  | (10 ~ 100) mA | 7.0 × 10 -4 | / KCTI-CQI-40311
|  | (0.1 ~ 1) A | 7.0 × 10 -4
|  | (1 ~ 10) A | 9.3 × 10 -4
|  | (10 ~ 20) A | 1.7 × 10 -3
교류전력 |  | (240V, 50 Hz)
|  | (0.24 ~ 1.2) W | 2.1 × 10 -2
|  | (1.2 ~ 2.4) W | 2.1 × 10 -2
|  | (2.4 ~ 6) W | 1.0 × 10 -2
|  | (6 ~ 12) W | 2.1 × 10 -2
|  | (12 ~ 24) W | 2.1 × 10 -2
|  | (24 ~ 60) W | 1.0 × 10 -2
|  | (60 ~ 120) W | 2.5 × 10 -2
|  | (120 ~ 240) W | 2.1 × 10 -2
|  | (240 ~ 600) W | 2.1 × 10 -2
|  | (0.6 ~ 1.2) kW | 1.0 × 10 -2
|  | (1.2 ~ 2.4) kW | 5.2 × 10 -3
|  | (2.4 ~ 4.8) kW | 2.7 × 10 -3
(240V, 60 Hz)
|  | (0.24 ~ 1.2) W | 2.1 × 10 -2
|  | (1.2 ~ 2.4) W | 2.1 × 10 -2
|  | (2.4 ~ 6) W | 1.0 × 10 -2
|  | (6 ~ 12) W | 2.1 × 10 -2
|  | (12 ~ 24) W | 2.1 × 10 -2
|  | (24 ~ 60) W | 1.0 × 10 -2
|  | (60 ~ 120) W | 2.5 × 10 -2
|  | (120 ~ 240) W | 2.1 × 10 -2
|  | (240 ~ 600) W | 2.1 × 10 -2
|  | (0.6 ~ 1.2) kW | 1.0 × 10 -2
|  | (1.2 ~ 2.4) kW | 5.2 × 10 -3
|  | (2.4 ~ 4.8) kW | 3.1 × 10 -3
(120V, 50 Hz)
|  | (0.12 ~ 0.6) W | 2.2 × 10 -2
|  | (0.6 ~ 1.2) W | 2.5 × 10 -2
|  | (1.2 ~ 3) W | 1.3 × 10 -2
|  | (3 ~ 6) W | 2.2 × 10 -2
|  | (6 ~ 12) W | 2.5 × 10 -2
|  | (12 ~ 30) W | 1.3 × 10 -2
|  | (30 ~ 60) W | 2.2 × 10 -2
|  | (60 ~ 120) W | 2.5 × 10 -2
|  | (120 ~ 300) W | 2.1 × 10 -2
|  | (300 ~ 600) W | 1.0 × 10 -2
|  | (0.6 ~ 1.2) kW | 4.2 × 10 -3
|  | (1.2 ~ 2.4) kW | 3.0 × 10 -3
403. 교류 및 교류전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
교류 전력계, 피상.고조파 및 | 40311
무효전력계 등
교류전력 |  | (120V, 60 Hz) |  | Meter Calibrator
|  | (0.12 ~ 0.6) W | 2.2 × 10 -2 | / KCTI-CQI-40311
|  | (0.6 ~ 1.2) W | 2.5 × 10 -2
|  | (1.2 ~ 3) W | 1.3 × 10 -2
|  | (3 ~ 6) W | 2.2 × 10 -2
|  | (6 ~ 12) W | 2.5 × 10 -2
|  | (12 ~ 30) W | 1.3 × 10 -2
|  | (30 ~ 60) W | 2.2 × 10 -2
|  | (60 ~ 120) W | 2.5 × 10 -2
|  | (120 ~ 300) W | 2.1 × 10 -2
|  | (300 ~ 600) W | 1.0 × 10 -2
|  | (0.6 ~ 1.2) kW | 4.2 × 10 -3
|  | (1.2 ~ 2.4) kW | 3.0 × 10 -3
역률 |  | 50 Hz, 60 Hz
(LEAD & LAG) |  | 1 | 0.001
|  | 0.9 | 0.002
|  | 0.8 | 0.002
|  | 0.7 | 0.003
|  | 0.6 | 0.003
|  | 0.5 | 0.004
|  | 0.4 | 0.005
|  | 0.3 | 0.007
|  | 0.2 | 0.01
|  | 0.1 | 0.021
교류 전원공급기 | 40312
교류전압 |  | 40 Hz ~ 100 Hz |  | Digital Multimeter
|  | 1 V | 0.058 V | Frequency Counter
|  | (1 ~ 10) V | 5.8 × 10 -3 | AC Load
|  | (10 ~ 100) V | 6.0 × 10 -4 | Current Shunt
|  | (100 V ~ 1 000) V | 1.6 × 10 -4 | / KCTI-CQI-40312
100 Hz ~ 1 kHz
|  | 1 V | 0.058 V
|  | (1 ~ 10) V | 5.8 × 10 -3
|  | (10 ~ 100) V | 5.9 × 10 -4
|  | (100 V ~ 1 000) V | 1.6 × 10 -4
교류전류 |  | 40 Hz ~ 1 kHz
|  | 100 mA | 0.14 mA
|  | (0.1 ~ 1) A | 1.4 × 10 -3
|  | (1 ~ 10) A | 3.2 × 10 -3
|  | (10 ~ 20) A | 6.2 × 10 -3
403. 교류 및 교류전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
내전압/전기 안전 시험기 | 40313
직류전압 |  | (0 ~ 0.5) kV | 1.0 × 10 -2 | Decade Resistance Box
|  | (0.5 ~ 1) kV | 2.0 × 10 -3 | Digital Multimeter
|  | (1 ~ 5) kV | 7.0 × 10 -3 | High Voltage Meter
|  | (5 ~ 10) kV | 6.4 × 10 -3 | / KCTI-CQI-40313
|  | (10 ~ 20) kV | 1.8 × 10 -2
|  | (20 ~ 50) kV | 2.0 × 10 -2
|  | (50 ~ 100) kV | 1.9 × 10 -2
교류전압 |  | 50 Hz
|  | (0.1 ~ 0.5) kV | 1.0 × 10 -2
|  | (0.5 ~ 1) kV | 2.0 × 10 -3
|  | (1 ~ 5) kV | 7.0 × 10 -3
|  | (5 ~ 10) kV | 6.4 × 10 -3
60 Hz
|  | (0.1 ~ 0.5) kV | 1.0 × 10 -2
|  | (0.5 ~ 1) kV | 2.0 × 10 -3
|  | (1 ~ 5) kV | 7.0 × 10 -3
|  | (5 ~ 10) kV | 6.4 × 10 -3
|  | (10 ~ 20) kV | 1.9 × 10 -2
|  | (20 ~ 50) kV | 2.2 × 10 -2
|  | (50 ~ 100) kV | 2.0 × 10 -2
직류전류 |  | (0 ~ 0.5) mA | 2.0 × 10 -3
|  | (0.5 ~ 1) mA | 1.0 × 10 -3
|  | (1 ~ 2) mA | 5.0 × 10 -4
|  | (2 ~ 5) mA | 2.0 × 10 -4
|  | (5 ~ 10) mA | 1.0 × 10 -4
|  | (10 ~ 20) mA | 2.5 × 10 -4
|  | (20 ~ 50) mA | 1.2 × 10 -4
|  | (50 ~ 100) mA | 7.0 × 10 -5
교류전류 |  | 50 Hz
|  | (0 ~ 0.5) mA | 1.5 × 10 -3
|  | (0.5 ~ 1) mA | 1.0 × 10 -3
|  | (1 ~ 2) mA | 2.0 × 10 -3
|  | (2 ~ 5) mA | 1.0 × 10 -3
|  | (5 ~ 10) mA | 6.0 × 10 -4
|  | (10 ~ 20) mA | 1.8 × 10 -3
|  | (20 ~ 50) mA | 8.6 × 10 -4
|  | (50 ~ 100) mA | 5.6 × 10 -4
60 Hz
|  | (0 ~ 0.5) mA | 2.0 × 10 -3
|  | (0.5 ~ 1) mA | 1.0 × 10 -3
|  | (1 ~ 2) mA | 2.0 × 10 -3
|  | (2 ~ 5) mA | 1.0 × 10 -3
|  | (5 ~ 10) mA | 6.0 × 10 -4
|  | (10 ~ 20) mA | 1.8 × 10 -3
|  | (20 ~ 50) mA | 8.6 × 10 -4
|  | (50 ~ 100) mA | 5.6 × 10 -4
403. 교류 및 교류전력
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
교류전압계/전위차, 실효치 | 40318
교류전압 |  | 40 Hz
|  | (10 ~ 100) mV | 1.8 × 10 -3 | Meter Calibrator
|  | (0.1 ~ 1) V | 4.6 × 10 -4 | / KCTI-CQI-40318
50 Hz ~ 1 kHz
|  | (1 ~ 10) mV | 4.4 × 10 -4
|  | (10 ~ 100) mV | 1.1 × 10 -3
|  | (0.1 ~ 1) V | 2.8 × 10 -4
|  | (1 ~ 100) V | 2.6 × 10 -4
|  | (100 ~ 1 000) V | 3.7 × 10 -4
404. 기타직류 및 저주파
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
전원 주파수계 | 40410
주파수 |  | 45 Hz | 2.6 × 10 -4 | Meter Calibrator
|  | (0.045 ~ 10) kHz | 5.9 × 10 -4 | / KCTI-CQI-40410
누설전류 시험기 | 40416
직류전류 |  | 10 μA | 2.5 × 10 -3 | Meter Calibrator
|  | (10 ~ 100) μA | 7.1 × 10 -4 | / KCTI-CQI-40416
|  | (0.1 ~ 100) mA | 6.1 × 10 -4
교류전류 |  | 50 Hz ~ 1 kHz
|  | 10 μA | 1.2 × 10 -2
|  | (10 ~ 100) μA | 1.4 × 10 -3
|  | (0.1 ~ 100) mA | 6.3 × 10 -4
교류전압 |  | 50 Hz ~ 1 kHz
|  | 10 mV | 1.1 × 10 -3
|  | 100 mV | 6.4 × 10 -4
|  | (0.1 ~ 100) V | 6.3 × 10 -4
404. 기타직류 및 저주파
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
직/교류 전자부하 | 40417
직류전압 |  | 100 mV | 6.7 × 10 -5 | Meter Calibrator
|  | (0.1 ~ 1 000) V | 6.2 × 10 -5 | / KCTI-CQI-40417
직류전류 |  | 10 μA | 2.5 × 10 -3
|  | (0.01 ~ 100) mA | 4.1 × 10 -4
|  | (0.1 ~ 1) A | 2.8 × 10 -4
|  | (1 ~ 10) A | 6.4 × 10 -4
|  | (10 ~ 100) A | 1.6 × 10 -4
(50 ~ 400 )Hz
교류전압 |  | 100 mV | 2.8 × 10 -4
|  | (0.1 ~ 100) V | 2.5 × 10 -4
|  | (100 ~ 1 000) V | 4.0 × 10 -4
(50 ~ 400 )Hz
교류전류 |  | 10 μA | 1.7 × 10 -2
|  | (0.01 ~ 100) mA | 1.8 × 10 -3
|  | (0.1 ~ 1) A | 1.4 × 10 -4
|  | (1 ~ 10) A | 2.5 × 10 -4
아날로그/디지털 멀티미터; | 40419
하부 속성 개별 인정
직류 전압 |  | (10 ~ 100) mV | 3.7 × 10 -5 | Meter Calibrator
|  | (0.1 ~ 1) V | 1.6 × 10 -5 | / KCTI-CQI-40419
|  | (1 ~ 10) V | 1.7 × 10 -5
|  | (10 ~ 100) V | 2.3 × 10 -5
|  | (100 ~ 1 000) V | 2.3 × 10 -5
교류 전압 |  | 40 Hz ~ 1 kHz
|  | (1 ~ 10) mV | 1.6 × 10 -3
|  | (10 ~ 100) mV | 4.5 × 10 -4
|  | (0.1 ~ 1) V | 4.1 × 10 -4
|  | (1 ~ 10) V | 4.3 × 10 -4
1 kHz ~ 20 kHz
|  | (1 ~ 10) mV | 9.4 × 10 -4
|  | (10 ~ 100) mV | 2.9 × 10 -4
|  | (0.1 ~ 1) V | 2.9 × 10 -4
|  | (1 ~ 10) V | 3.5 × 10 -4
20 kHz ~ 100 kHz
|  | (1 ~ 10) mV | 5.4 × 10 -3
|  | (10 ~ 100) mV | 1.3 × 10 -3
|  | (0.1 ~ 1) V | 9.7 × 10 -4
|  | (1 ~ 10) V | 1.2 × 10 -3
50 Hz ~ 1 kHz
|  | (1 ~ 100) V | 2.5 × 10 -4
|  | (100 ~ 1 000) V | 3.6 × 10 -4
1 kHz ~ 20 kHz
|  | (10 ~ 100) V | 3.8 × 10 -4
404. 기타직류 및 저주파
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
아날로그/디지털 멀티미터; 하부 속성 개별 인정 | 40419
직류 전류 |  | (0.1 ~ 1) mA | 1.7 × 10 -4 | Meter Calibrator
|  | (1 ~ 10) mA | 1.4 × 10 -4 | / KCTI-CQI-40419
|  | (10 ~ 100) mA | 1.4 × 10 -4
|  | (0.1 ~ 1) A | 2.8 × 10 -4
|  | (1 ~ 10) A | 6.4 × 10 -4
교류 전류 |  | 100 Hz ~ 1 kHz
|  | (0.1 ~ 1) mA | 2.3 × 10 -4
|  | (1 ~ 10) mA | 2.8 × 10 -4
|  | (10 ~ 100) mA | 2.9 × 10 -4
|  | (0.1 ~ 1) A | 2.6 × 10 -4
50 Hz ~ 1 kHz
|  | (1 ~ 10) A | 3.9 × 10 -4
저항 |  | (1 ~ 10) Ω | 1.1 × 10 -4
|  | (10 ~ 100) Ω | 1.1 × 10 -4
|  | (0.1 ~ 1) kΩ | 5.1 × 10 -5
|  | (1 ~ 10) kΩ | 5.1 × 10 -5
|  | (10 ~ 100) kΩ | 6.0 × 10 -5
|  | (0.1 ~ 1) MΩ | 5.4 × 10 -5
|  | (1 ~ 10) MΩ | 1.6 × 10 -4
|  | (10 ~ 100) MΩ | 6.4 × 10 -4
파형측정기 | 40421
직류전압 |  | 1 mV | 4.7 × 10 -2 | Oscilloscope Calibrator
|  | (1 ~ 2) mV | 2.4 × 10 -2 | Frequency Counter
|  | (2 ~ 5) mV | 1.0 × 10 -2 | Digital Multimeter
|  | (5 ~ 10) mV | 5.2 × 10 -3 | Oscilloscope
|  | (10 ~ 20) mV | 2.9 × 10 -3 | / KCTI-CQI-40421
|  | (20 ~ 50) mV | 1.6 × 10 -3
|  | (50 ~ 100) mV | 1.1 × 10 -3
|  | (0.1 ~ 0.2) V | 1.0 × 10 -3
|  | (0.2 ~ 0.5) V | 8.0 × 10 -4
|  | (0.5 ~ 1) V | 7.0 × 10 -4
|  | (1 ~ 2) V | 6.5 × 10 -4
|  | (2 ~ 5) V | 6.0 × 10 -4
|  | (5 ~ 10) V | 6.0 × 10 -4
|  | (10 ~ 20) V | 1.0 × 10 -3
|  | (20 ~ 50) V | 8.0 × 10 -4
|  | (50 ~ 100) V | 6.0 × 10 -4
교류전압 |  | 2 mV | 2.5 × 10 -2
|  | (2 ~ 5) mV | 1.1 × 10 -2
|  | (5 ~ 10) mV | 5.8 × 10 -3
|  | (10 ~ 20) mV | 3.5 × 10 -3
|  | (20 ~ 50) mV | 2.2 × 10 -3
|  | (50 ~ 100) mV | 1.7 × 10 -3
|  | (0.1 ~ 0.2) V | 1.5 × 10 -3
|  | (0.2 ~ 0.5) V | 1.4 × 10 -3
|  | (0.5 ~ 1) V | 1.3 × 10 -3
|  | (1 ~ 2) V | 1.2 × 10 -3
|  | (2 ~ 5) V | 1.2 × 10 -3
|  | (5 ~ 10) V | 1.2 × 10 -3
|  | (10 ~ 20) V | 1.5 × 10 -3
|  | (20 ~ 50) V | 1.2 × 10 -3
|  | (50 ~ 100) V | 1.2 × 10 -3
404. 기타직류 및 저주파
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
파형측정기 | 40421
주기 |  | 5 ns | 2.0 × 10 -4 | Oscilloscope Calibrator
|  | (5 ~ 10) ns | 1.0 × 10 -3 | Frequency Counter
|  | (10 ~ 20) ns | 5.0 × 10 -4 | Digital Multimeter
|  | (20 ~ 50) ns | 2.0 × 10 -4 | Oscilloscope
|  | (50 ~ 100) ns | 1.0 × 10 -3 | / KCTI-CQI-40421
|  | (100 ~ 200) ns | 5.0 × 10 -4
|  | (200 ~ 500) ns | 2.0 × 10 -4
|  | (0.5 ~ 1) μs | 1.0 × 10 -3
|  | (1 ~ 2) μs | 5.0 × 10 -4
|  | (2 ~ 5) μs | 2.0 × 10 -4
|  | (5 ~ 10) μs | 1.0 × 10 -3
|  | (10 ~ 20) μs | 5.0 × 10 -4
|  | (20 ~ 50) μs | 2.0 × 10 -4
|  | (50 ~ 100) μs | 1.0 × 10 -3
|  | (100 ~ 200) μs | 5.0 × 10 -4
|  | (200 ~ 500) μs | 2.0 × 10 -4
|  | (0.5 ~ 1) ms | 1.0 × 10 -3
|  | (1 ~ 2) ms | 5.0 × 10 -4
|  | (2 ~ 5) ms | 2.0 × 10 -4
|  | (5 ~ 10) ms | 1.0 × 10 -3
|  | (10 ~ 20) ms | 5.0 × 10 -4
|  | (20 ~ 50) ms | 2.0 × 10 -4
|  | (50 ~ 100) ms | 1.0 × 10 -3
|  | (100 ~ 200) ms | 5.0 × 10 -4
|  | (200 ~ 500) ms | 8.0 × 10 -4
|  | (0.5 ~ 1) s | 2.0 × 10 -3
|  | (1 ~ 2) s | 2.5 × 10 -3
|  | (2 ~ 5) s | 6.0 × 10 -3
정현파 레벨 |  | 50 kHz | 1.7 × 10 -2
Ref. 50 kHz |  | (50 ~ 100) kHz | 1.7 × 10 -2 | Oscilloscope Calibrator
|  | (0.1 ~ 1) MHz | 1.7 × 10 -2 | Frequency Counter
|  | (1 ~ 10) MHz | 2.9 × 10 -2 | Digital Multimeter
|  | (10 ~ 50) MHz | 2.9 × 10 -2 | Oscilloscope
|  | (50 ~ 100) MHz | 2.9 × 10 -2 | / KCTI-CQI-40421
|  | (100 ~ 300) MHz | 2.9 × 10 -2
|  | (300 ~ 500) MHz | 2.9 × 10 -2
|  | (500 ~ 600) MHz | 3.7 × 10 -2
|  | (600 ~ 1 000) MHz | 3.7 × 10 -2
|  | (1 000 ~ 1 300) MHz | 4.7 × 10 -2
|  | (1 300 ~ 1 500) MHz | 4.7 × 10 -2
|  | (1 500 ~ 1 800) MHz | 4.7 × 10 -2
|  | (1 800 ~ 2 000) MHz | 4.7 × 10 -2
|  | (2 000 ~ 2 100) MHz | 4.7 × 10 -2
501. 접촉식 온도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
온도 발생장치 ; | 50101
오븐, 전기로, 액체항온조,
빙점조, 드라이블럭교정기 등
오븐 |  | (-40 ~ 200) ℃ | 0.9 ℃ | 표준 온도계,백금저항식
온도 브릿지
전기로 |  | (200 ~ 1 100) ℃ | 1.0 ℃ | 온도 측정기
/ KCTI-CQI-50101
액체 항온조 |  | (-80 ~ 200) ℃ | 0.06 ℃
드라이블럭 교정기 |  | (-40 ~ 200) ℃ | 0.07 ℃
|  | (200 ~ 1 100) ℃ | 1.0 ℃
501. 접촉식 온도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
온도 지시계 ; 지시/기록/조절계 | 50102
온도 교정기 등
저항식(센서포함) |  | (-80 ~ -39) ℃ | 0.07 ℃ | 표준 온도계,백금저항식
|  | (-39 ~ 200) ℃ | 0.08 ℃ | 온도 브릿지,온도 측정기
전기로,Bath
저항식(센서제외) |  | (-80 ~ -39) ℃ | 0.15 ℃ | / KCTI-CQI-50102
|  | (-39 ~ 200) ℃ | 0.3 ℃
열전식(센서포함) |  | (-80 ~ -39) ℃ | 0.15 ℃
|  | (-39 ~ 200) ℃ | 0.2 ℃
|  | (200 ~ 1 100) ℃ | 1.3 ℃
열전식(센서제외)
R-Type |  | (0 ~ 1 100) ℃ | 1.1 ℃
S-Type |  | (0 ~ 1 100) ℃ | 1.1 ℃
B-Type |  | (600 ~ 1 100) ℃ | 1.1
K-Type |  | (-80 ~ -39) ℃ | 0.5 ℃
|  | (-39 ~ 1 100) ℃ | 0.5
N-Type |  | (0 ~ 1 100) ℃ | 0.6
E-Type |  | (-80 ~ -39) ℃ | 0.6 ℃
|  | (-39 ~ 700) ℃ | 0.6
J-Type |  | (-80 ~ -39) ℃ | 0.4 ℃
|  | (-39 ~ 600) ℃ | 0.4
T-Type |  | (-80 ~ -39) ℃ | 0.5 ℃
|  | (-39 ~ 300) ℃ | 0.5
유리제 온도계 ; 유리제 온도계, | 50103
벡크만 온도계 등
유리제 온도계 |  | (-39 ~ 200) ℃ | 0.04 ℃ | 백금저항식 온도 브릿지
/ KCTI-CQI-50103
저항식 온도계 ; | 50104
백금저항 온도계,
측온저항체, 써미스터 등
|  | (-80 ~ 200) ℃ | 0.06 ℃ | 백금저항식 온도 브릿지
산업용 저항온도계 |  |  |  | / KCTI-CQI-50104
열팽창식 온도계 ; | 50105
바이메탈 온도계, 기체또는
액체 충만식 온도계 등
바이메탈 온도계 |  | (-39 ~ 200) ℃ | 0.3 ℃ | 백금저항식 온도 브릿지
/ KCTI-CQI-50105
열전대 ; 귀금속, 비금속, | 50106
순금속, 특수 등
비금속 열전대 |  | (-80 ~ -39) ℃ | 0.06 ℃ | 표준 열전대
|  | (-39 ~ 200) ℃ | 0.3 ℃ | / KCTI-CQI-50106
|  | (200 ~ 1 100) ℃ | 1.3 ℃
귀금속 열전대 |  | (0 ~ 200) ℃ | 0.3 ℃ | 표준 열전대
|  | (200 ~ 1 100) ℃ | 1.6 ℃ | / KCTI-CQI-50106
온도 변환기 | 50107 | (-39 ~ 200) ℃ | 0.3 ℃ | 표준 온도계,백금저항식
|  | (200 ~ 1 100) ℃ | 1.2 ℃ | 온도 브릿지,전기로
/ KCTI-CQI-50107
502. 비접촉식 온도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
복사온도계 | 50204
|  | (0 ~ 50) ℃ | 0.6 ℃ | 기준복사 온도계
|  | (50 ~ 200) ℃ | 0.8 ℃ | / KCTI-CQI-50204
|  | (200 ~ 500) ℃ | 1.0 ℃
|  | (500 ~ 800) ℃ | 4.5 ℃
|  | (800 ~ 1 000) ℃ | 4.9 ℃
503. 습도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
상대습도 습도계; | 50302
고분자박막,모발 등
고분자박막,디지털 온습도계
습도 |  | (20 ~ 40) % R.H. | 2.1 % R.H.
|  | (40 ~ 60) % R.H. | 2.9 % R.H.
|  | (60 ~ 80) % R.H. | 3.2 % R.H.
|  | (80 ~ 90) % R.H. | 3.8 % R.H.
온도 |  | (-40 ~ 20) ℃ | 1.1 ℃ | 자동 노점계,항온항습기
|  | (20 ~ 90) ℃ | 1.4 ℃ | / KCTI-CQI-50302
모발 습도계
습도 |  | (20 ~ 90) % R.H. | 4.4 % R.H. | / KCTI-CQI-50302
온도 |  | (10 ~ 90) ℃ | 1.9 ℃ | 자동 노점계,항온항습기
건습구 습도계; 아스만 통풍, | 50303
저항온도계식 등
저항온도계식 등
습도 |  | (20 ~ 90) % R.H. | 3.9 % R.H. | 자동 노점계,항온항습기
온도 |  | (10 ~ 90) ℃ | 1.2 ℃ | / KCTI-CQI-50303
온·습도 기록계; | 50304
자기온습도기록계 등
습도 |  | (20 ~ 90) % R.H. | 3.7 % R.H. | / KCTI-CQI-50304
온도 |  | (-20 ~ 50) ℃ | 2.3 ℃ | 자동 노점계,항온항습기
노점/상대습도 변환기 | 50305
습도 |  | (20 ~ 40) % R.H. | 2.5 % R.H. | 자동 노점계,항온항습기
|  | (40 ~ 60) % R.H. | 2.7 % R.H. | / KCTI-CQI-50305
|  | (60 ~ 80) % R.H. | 3.4 % R.H.
|  | (80 ~ 90) % R.H. | 4.0 % R.H.
습도 발생장치; 이압력식/ | 50306
이온도식/분류식 습도발생
장치,항온항습기 등
습도 |  | (20 ~ 40) % R.H. | 2.5 % R.H. | 자동 노점계
|  | (40 ~ 60) % R.H. | 3.5 % R.H. | / KCTI-CQI-50306
|  | (60 ~ 80) % R.H. | 4.4 % R.H.
|  | (80 ~ 95) % R.H. | 5.0 % R.H.
온도 |  | (-40 ~ 0) ℃ | 0.8 ℃ | 온도기록계,PRT센서
|  | (0 ~ 90) ℃ | 1.0 ℃ | / KCTI-CQI-50306
601. 음향
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
소음계 | 60106 | 125 Hz | 0.3 dB | 다기능 음압 교정기
|  | 250 Hz | 0.2 dB | / KCTI-CQI-60106
|  | 500 Hz | 0.2 dB
|  | 1 kHz | 0.2 dB
|  | 2 kHz | 0.2 dB
|  | 4 kHz | 0.2 dB
|  | 8 kHz | 0.5 dB
701. 광도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
광 조도계 | 70101 | 0.5 lx | 3.1 × 10 -2 | 기준 광 조도계
|  | (0.5 ~ 50) lx | 2.5 × 10 -2 | / KCTI-CQI-70101
|  | (50 ~ 1 500) lx | 1.4 × 10 -2
|  | (1 500 ~ 5 000) lx | 1.8 × 10 -2

## 표지(영문)-시흥
SCOPE OF ACCREDITATION TO ISO/IEC 17025-2017 & KS Q ISO/IEC 17025-2017
Korea Calibration Technology Institute Co.,Ltd
206ho Ga-dong Hwasin office 17, Seochonsangga 4-gil, Siheung-city Gyeonggi-do, Republic of Korea
Phone No : 031)497-4694, Fax : 031)434-4694, e-mail : kcti@naver.com
CALIBRATION
Valid To : Mar. 19. 2027 | [7 empty] | Accreditation No. : KC07-212
In recognition of the successful completion of the KOLAS evaluation process,
accreditation is granted to this laboratory to perform the following calibrations
Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site
102. Linear dimension |  |  | 10517 | Stylus type roughness | Y | 20413 | Dial type vacuum gauges | Y
10201 | Balls | N |  | testers |  | 206. Volume
10206 | Dial/cylinder gauge testers | N | 10525 | Thread plug gauges | N | 20601 | Volumetric glassware | N
10209 | End bars | N | 10526 | Taper thread plug gauges | N | 20605 | Concrete air content meter | N
10210 | Extensometers, linear | Y | 10527 | Thread ring gauges | N | 20606 | Piston type volume meters | N
| displacement transducers |  | 10529 | V-blocks, box blocks | N | 207. Density
10211 | Filler gauges | N | 106. Various dimensional |  |  | 20702 | Liquid density meters | N
10213 | Gap gauges | N | 10601 | Inside/outside/gear tooth | Y | 20704 | Salinity meters | N
10216 | Height gauges/measuring | Y |  | calipers,caliper gauges |  | 20705 | Sucrose meters | N
| machines |  | 10603 | Cylinder/bore gauges | Y | 20707 | Chloride meters | N
10220 | Standard measuring machines | Y | 10604 | Depth gauges, depth | Y | 208. Viscosity
10223 | Electronic micrometers | N |  | micrometers |  | 20801 | (Kinematic viscometers; | N
10224 | Height micrometers, | N | 10605 | Dial/digital gauges | Y |  | capillary, etc.)
| riser blocks |  | 10609 | Micro indicators, Test | Y | 20802 | (Dynamic viscometers; | N
10227 | Standard tape rules, | N |  | indicators |  |  | rotaional, etc.)
| peripheral gauges |  | 10610 | Micrometer heads | N | 210. Hardness
10228 | Cylindrical plug/pin gauges | N | 10611 | 3-point micrometers | Y | 21001 | Brinell hardness testers | Y
| thread measuring wire gauges |  | 10612 | Inside micrometers | Y | 21002 | Rockwell hardness testers | Y
10229 | Radius gauges | N | 10613 | Outside micrometers | Y | 21003 | Shore hardness testers | Y
10230 | Cylindrical ring gauges | N | 10617 | Standard sieves | N | 21004 | Vickers hardness testers | Y
10232 | Step gauges | N | 10620 | Welding gauges | N | 21005 | Durometer hardness testers | N
10233 | Taper thickness gauges | N | 201. Mass |  |  | 21006 | Leeb hardness testers | Y
10234 | Ultrasonic thickness gauges | Y | 20102 | Auto-hopper scale balances | Y | 301. Time/frequency
10235 | Ultrasonic/coating | N | 20105 | Counter beam balances | Y | 30104 | Frequency meters/counters | N
| thickness specimens |  | 20109 | Electric balances | Y | 30106 | Time interval meters/ | Y
10236 | Coating thickness testers | Y | 20112 | Platform scale balances | Y |  | stop watches/timers
104. Form |  |  | 20113 | Spring scale balances | Y | 302. Velocity & revolution
10401 | Form testers | Y | 20116 | Weights | Y | 30201 | Standard RPM generators | Y
10406 | Parallel blocks | N | 202. Force |  |  | 30202 | Contact type tachometers | N
10407 | Precision surface plates | Y | 20203 | Tension/compression testing | Y | 30203 | Photo tachometers/strobo | N
10409 | Roundness measurement | Y |  | machines |  |  | scopes
| instruments |  | 20204 | Push-pull gauges | Y | 401. DC volatage & current
10413 | Straight rules | N | 203. Torque |  |  | 40101 | DC ammeters | Y
105. Complex geometry |  |  | 20302 | Torque measuring devices | Y | 40103 | DC voltage / current | Y
10502 | Bench centers | N | 20303 | Torque wrenches/drivers | Y |  | calibrators
10503 | Contact coordinate | Y | 204. Pressure |  |  | 40104 | Electrical temperature | N
| measuring machines |  | 20406 | Absolute pressure gauges | N |  | calibrators
10504 | Non-contact coordinate | Y | 20408 | Compound pressure gauges | Y | 40105 | DC current shunts | Y
| measuring machines |  | 20409 | Differential pressure gauges | Y | 40108 | DC power supplies | Y
10511 | Measuring microscopes, | Y | 20411 | Gauge pressure gauges | Y | 40112 | DC voltmeters | Y
| profile projectors |  | 20412 | Pressure transducers/ | Y | 402. Resistance, capacitance and inductance
10512 | Micro measuring microscopes | N |  | transmitters |  | 40205 | Earth testers | Y
Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site
40210 | Insulation testers | Y | 50105 | Thermal expansion | Y
40214 | Resistance meters | Y |  | thermometers;
40215 | Resistor | Y |  | bimetal, gas or liquid type
40217 | Impedance bridges/ | Y | 50106 | Thermomecouples; | Y
| LCR meters |  |  | noble metal, base metal,
403. AC voltage, current & power |  |  |  | pure metal special type, ect.
40301 | AC ammeters | Y | 50107 | Temperature transducers | Y
40302 | Clamp ammeters/voltmeters | Y | 502. Non contact thermometry
40303 | AC voltage/current | Y | 50204 | Standard radiation | N
| calibrators |  |  | hermometers
40305 | AC current shunts | Y | 503. Humidity
40310 | Power factor meter | Y | 50302 | Relative humidity | N
40311 | AC power meters | Y |  | hygrometers; polimer
40312 | AC power supplies | Y |  | thin film, hair, etc.
40313 | Puncture/safety testers | Y | 50303 | Psychrometers; | N
40318 | AC voltmeter | Y |  | Assmann ventilated,
404. Other DC & LF measurements |  |  |  | PRT type, etc.
40410 | Line frequency meters | Y | 50304 | Temperature humidity | N
40416 | Leakage current testers | Y |  | recorders; hygrothermo-
40417 | Electronic AC/DC loads | Y |  | graph, etc.
40419 | Analogue/digital multimeters | Y | 50305 | Transducers; | N
40421 | Oscilloscopes | Y |  | dew-point/relative humidity
501. Contact thermometry |  |  | 50306 | Humidity generators; | Y
50101 | Temperature generators; | Y |  | two-pressure,
| ovens, furnaces, |  |  | two-temperature,
| isothermal liquid baths, |  |  | flow mixing humidity
| ice-point baths, |  |  | gererator,
| dry-block calibrators |  |  | constant temperature and
50102 | Temperature indicators | Y |  | humidity chamber, etc.
| /recorders/ |  | 601. Sound in air
| controllers, temperature |  | 60106 | Sound level meters | N
| calibrators |  | 701. Photometry
50103 | Glass thermometers; | N | 70101 | Iluminance meters | N
liquid-in-glass, Beckmann
50104 | Resistance thermometers; | Y
SPRT, IPRT, thermistors,
etc.
Note
1. This laboratory provides calibration services in permanent standard laboratory and at on-site.
2. Laboratory conducts on-site calibration should meet requirements of KOLAS-SR-008.
3. On-site calibration is allowed to items with marking 'Y', not allowed to items with marking 'N'.
4. Calibration and Measurement Capability (CMC) means capabilities provided by accredited calibration laboratories. It expresses the lowest uncertainty of measurement that can be achieved during a calibration. CMC normally is quoted as an expanded uncertainty at a coverage probability of 95 %, which usually requires the use of a coverage factor of k=2.
5. Due to the calibration environment such as reference standards or customers’ facilities, it is note that uncertainty of measurement on a calibration certificate may be expressed larger than CMC on scope of accreditation in general.

## 교정측정능력(영문)-시흥
102. Linear dimension
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Balls | 10201 | (5 ~ 50) mm | 1.02 + (0.0032 × l02) μm | Standard measuring machines
|  |  | (l0:mm) | / KCTI-CQI-10201
Dial/cylinder gauge testers | 10206 | (0 ~ 100) mm | 0.372 + (0.0022 × l02) μm | Electronic micrometer
|  |  | (l0:mm) | Gauge blocks
/ KCTI-CQI-10206
End bars | 10209 | (25 ~ 500) mm | 1.22 + (0.0072 × l02) μm | Electronic micrometer
|  |  | (l0:mm) | Gauge blocks
/ KCTI-CQI-10209
Extensometers,linear | 10210 | (0 ~ 600) mm | 1.32 + (0.004 22 × l02) μm | Electronic micrometer
displacement transducers |  |  | (l0:mm) | Gauge blocks
/ KCTI-CQI-10210
Filler gauges | 10211 | (0.01 ~ 3) mm | 1.1 μm | Micrometers,Bench
/ KCTI-CQI-10211
Gap gauges | 10213 | (5 ~ 200) mm | 3.12 + (0.0062 × l02) μm | Electronic micrometer
|  |  | (l0:mm) | / KCTI-CQI-10213
Height gauges/measuring | 10216 | (0 ~ 1 000) mm | 1.62 + (0.0032 × l02) μm | Gauge blocks,Step gauge
machines |  |  | (l0:mm) | / KCTI-CQI-10216
Standard measuring machines | 10220 | (0 ~ 500) mm | 0.32 + (0.0032 × l02) μm | Gauge blocks,
|  |  | (l0:mm) | Long Gauge blocks
/ KCTI-CQI-10220
Electronic micrometers | 10223 | (0 ~ 5) mm | 0.15 μm | Gauge blocks
/ KCTI-CQI-10223
Height micrometers,riser blocks | 10224
Height micrometers
Block |  | (0 ~ 600) mm | 1.32 + (0.005 22 × l02) mm | Electronic micrometer
|  |  | (l0:m) | Gauge blocks
Head |  | (0 ~ 25) mm | 1.42 + (0.0042 × l02) mm | / KCTI-CQI-10224
(l0:m)
Standard tape rules, | 10227
peripheral gauges |  | (0 ~ 5) m | 0.072 + (0.0122 × l02) mm | Linear scales
|  |  | (l0:m) | / KCTI-CQI-10227
|  | (5 ~ 50) m | 0.312 + (0.0122 × l02) mm
(l0:m)
Cylindrical plug/pin gauges, | 10228
thread measuring wire gauges
Cylindrical plug/pin gauges |  | (0.1 ~ 100) mm | 2.0 μm | Micrometers,Bench
Standard measuring machines
thread measuring wire gauges |  | (0.1 ~ 10) mm | 1.0 μm | / KCTI-CQI-10228
Radius gauges | 10229 | (0 ~ 50) mm | 4.1 μm | Measuring microscopes
/ KCTI-CQI-10229
Cylindrical ring gauges | 10230 | (3 ~ 150) mm | 1.5 μm | Standard measuring
/ KCTI-CQI-10230
Step gauges | 10232 | (0 ~ 1 000) mm | 1.42 + (0.0042 × l02) μm | Electronic micrometer
|  |  | (l0:mm) | Gauge blocks
/ KCTI-CQI-10232
102. Linear dimension
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Taper thickness gauges | 10233 | (0 ~ 150) mm | 8 μm | Measuring microscopes
/ KCTI-CQI-10233
Ultrasonic thickness gauges | 10234 | (0 ~ 100) mm | 8 μm | Gauge blocks
/ KCTI-CQI-10234
Ultrasonic/coating | 10235
thickness specimens
Coating |  | (1 ~ 5 000) μm | 1.5 μm | Standard measuring
machines
Ultrasonic thickness specimens |  | (0 ~ 200) mm | 1.02 + (0.0032 × l02) μm | / KCTI-CQI-10235
(l0:mm)
Coating thickness testers | 10236 | (0 ~ 1.6) mm | 2.5 μm | Standard thickness
/ KCTI-CQI-10236
104. Form
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Form testers | 10401
Vertical (Z) |  | (0 ~ 100) mm | 0.38 μm | Form standard thickness
Horizontal (X) |  | (1 ~ 50) mm | 0.91 μm | Gauge blocks
Angle |  | (0 ~ 90)˚ | 4.0˝ | Angle gauge blocks
/ KCTI-CQI-10401
Parallel blocks | 10406
Floor plan |  | (0 ~ 300) mm | 0.6 μm | Electronic micrometer
Parallelism |  | (0 ~ 300) mm | 0.6 μm | Gauge blocks
Differences in height and |  | (0 ~ 1) mm | 0.8 μm | / KCTI-CQI-10406
width of pairs
Precision surface plates | 10407
Diagonal length |  | (0 ~ 700) mm | 0.7 μm | Electrical levels
|  | (700 ~ 1 400) mm | 1.4 μm | / KCTI-CQI-10407
|  | (1 400 ~ 2 800) mm | 2.1 μm
|  | (2 800 ~ 4 000) mm | 2.6 μm
|  | (4 000 ~ 6 000) mm | 3.5 μm
Roundness measurement | 10409
instruments |  |  |  | Roundness standard
Detector accuracy |  | (0 ~ 300) μm | 0.40 μm | thickness
Spindle Rotation accuracy |  | (0 ~ 360)˚ | 20 nm | Roundness standard radius
Spindle shaft accuracy |  | (0 ~ 360)˚ | 40 nm | / KCTI-CQI-10409
Straight rules | 10413 |  |  | Linear scales
Scale spacing accuracy |  | (0 ~ 3 000) mm | 0.04 mm | /KCTI-CQI-10413
Straightness |  | (0 ~ 3 000) mm | 12 μm
105. Complex geometry
Measured Quantity Instrument or Gauge | Field code | Range | CMC (The Confidence Level is about 95 %) | Comments
Bench centers | 10502
Amount of parallelism Center |  | (100 ~ 300) mm | 3.4 μm | Electronic micrometer
Bed surface plan |  | (100 ~ 300) mm | 2.2 μm | Test Bar
Sheep Center height difference |  | (100 ~ 300) mm | 3.0 μm | / KCTI-CQI-10502
105. Complex geometry
Measured Quantity Instrument or Gauge | Field code | Range | CMC (The Confidence Level is about 95 %) | Comments
Contact coordinate measuring | 10503
machines
X,Y,Z accuracy |  | (0 ~ 1 000) mm | 0.82 + (0.0042 × l02) μm | Right angle rule
|  |  | (l0:mm) | Gauge blocks
Accuray |  | (0 ~ 1 000) mm | 0.92 + (0.0042 × l02) μm | Step gauge
|  |  | (l0:mm) | / KCTI-CQI-10503
Straightness |  | (0 ~ 500) mm | 3.0 μm
Right angle |  | (0 ~ 500) mm | 2.0˝
Non-contact coordinate | 10504
measuring machines
Accuracy instructions |  | (0 ~ 300) mm | 1.1 μm | Standard scales
|  | (300 ~ 500) mm | 1.7 μm | Right angle rule
Squareness(Y/X) |  | - | 5˝ | / KCTI-CQI-10504
Measuring microscopes, | 10511
Profile projectors
Measuring microscopes |  | (0 ~ 300) mm | 2.0 μm | Standard scales
|  | (300 ~ 500) mm | 3.1 μm | / KCTI-CQI-10511
Right angle |  | (0~200) mm | 3.1 μm
Profile projectors |  | (0 ~ 300) mm | 3.5 μm | Standard scales
Right angle |  | 100 mm | 2.6 μm
Magnification error |  | 10 magnification | 0.026 %
Projection plant rotation angle |  | (0 ~ 360)˚ | 2´
Micro measuring microscopes | 10512 | (0 ~ 1) mm | 1.0 μm | Standard scales
|  | (1 ~ 10) mm | 1.8 μm | / KCTI-CQI-10512
Stylus type roughness testers | 10517
Ra |  | (0 ~ 3.3) μm | 0.05 μm | Stylus Type Roughness
Rz |  | (0 ~ 13) μm | 0.18 μm | / KCTI-CQI-10517
RSm |  | (0 ~ 1) mm | 0.06 μm
Step height |  | (0 ~ 10) μm | 0.20 μm
Thread plug gauges | 10525
Pitch diameter |  | (1 ~ 50) mm | 1.9 μm | Micrometers,Bench
Pitch |  | (0 ~ 3) mm | 1.3 μm | wire gauges
Major diameter |  | (0 ~ 50) mm | 1.5 μm | / KCTI-CQI-10525
Half angle of thread |  | (0 ~ 60)˚ | 2.0´
Taper thread plug gauges | 10526
Smallend pitch diameter |  | (4 ~ 50) mm | 5.7 μm | Gauge blocks,Standard
Bigend pitch diameter |  | (4 ~ 50) mm | 5.7 μm | measuring machines
Smallend Major diameter |  | (4 ~ 50) mm | 5.7 μm | Electronic micrometer
Bigend Major diameter |  | (4 ~ 50) mm | 5.7 μm | Bench wire gauges
Pitch |  | (0.5 ~ 4) mm | 2.1 μm | highe master,Measuring
Half angle of thread |  | (0 ~ 45)˚ | 1´ 40˝ | microscopes
Half angle of taper |  | (0 ~ 45)˚ | 58˝ | / KCTI-CQI-10526
Gauge length |  | (0 ~ 100) mm | 2.0 μm
Thread ring gauges | 10527 |  |  | Standard measuring
Pitch diameter |  | (4 ~ 50) mm | 2.3 μm | machines,3-point
Inside diameter |  | (4 ~ 50) mm | 2.2 μm | Micrometers
/ KCTI-CQI-10527
105. Complex geometry
Measured Quantity Instrument or Gauge | Field code | Range | CMC (The Confidence Level is about 95 %) | Comments
V-blocks, box blocks | 10529
Floor plan of bottom |  | (0 ~ 150) mm | 1.8 μm | Electronic micrometer
V plan view of the surface |  | (0 ~ 150) mm | 1.8 μm | Test Bar
Parallelism of the circle |  | (0 ~ 150) mm | 3.8 μm | / KCTI-CQI-10529
passing through the above
bottom surface with V
The inclination of the V |  | (0 ~ 150) mm | 1.3 μm
groove on the bottom surface
Parallelism between the sides |  | (0 ~ 150) mm | 3.8 μm
and the V side upper circle
The mutual difference in the |  | (0 ~ 150) mm | 3.8 μm
height of the V-plane for one
pair of V blocks
106. Various dimensional
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Inside/outside/gear tooth | 10601
calipers,caliper gauges
Inside/Outside calipers |  | (0 ~ 1 500) mm | 102 + (0.0072 × l02) μm | Step gauges
|  |  | (l0:mm) | / KCTI-CQI-10601
Caliper gauges |  | (0 ~ 100) mm | 1.12 + (0.0042 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10601
Cylinder/bore gauges | 10603 | (0 ~ 1.2) mm | 0.7 μm | Dial gauge testers
/ KCTI-CQI-10603
Depth gauges,Depth micrometers | 10604
Depth gauges |  | (0 ~ 500) mm | 82 + (0.0042 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10604
Depth micrometers |  | (0 ~ 150) mm | 2.02 + (0.0072 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10604
Dial/Digital gauges | 10605 | (0 ~ 25) mm | 1.0 μm | Dial gauge testers
|  | (25 ~ 100) mm | 1.6 μm | / KCTI-CQI-10605
Micro indicators, Test | 10609
indicators
Micro indicators |  | (0 ~ 1) mm | 0.7 μm | Dial gauge testers
/ KCTI-CQI-10609
Test indicators |  | (0 ~ 1) mm | 0.7 μm | Dial gauge testers
/ KCTI-CQI-10609
Micrometer heads | 10610 | (0 ~ 100) mm | 0.82 + (0.0072 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10610
3-point micrometers | 10611 | (3 ~ 5) mm | 1.6 μm | Ring Gauge
|  | (5 ~ 100) mm | 2.3 μm | / KCTI-CQI-10611
|  | (100 ~ 150) mm | 2.5 μm
Inside micrometers | 10612
Inside micrometers |  | (5 ~ 200) mm | 1.22 + (0.0072 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10612
Micrometers, bar type |  | (50 ~ 500) mm | 1.32 + (0.0042 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10612
106. Various dimensional
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Outside micrometers | 10613
Outside micrometers |  | (0 ~ 25) mm | 0.52 + (0.0022 × l02) μm | Gauge blocks
(l0:mm)
|  | (25 ~ 125) mm | 1.02 + (0.0032 × l02) μm | / KCTI-CQI-10613
(l0:mm)
|  | (125 ~ 500) mm | 1.22 + (0.0052 × l02) μm
(l0:mm)
Micrometers, V-anvil |  | (2.5 ~ 25) mm | 1.3 μm | Pin Gauge
/ KCTI-CQI-10613
Standard sieves | 10617
Diameter of wire rod |  | (0.02 ~ 10) mm | 3.2 μm | Measuring microscopes
sieve scale |  | (0.02 ~ 100) mm | 4.8 μm | / KCTI-CQI-10617
Welding gauges | 10620
Length, height |  | (0 ~ 150) mm | 0.2 mm | Measuring microscopes
Neck thickness |  |  |  | / KCTI-CQI-10620
a graduated ruler
Taper thickness gauges
Angle |  | (0 ~ 90)˚ | 0.2˚
201. Mass
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Auto-hopper scale balances | 20102
|  | (0 ~ 10) kg | 1.5 g | Weight
|  | (10 ~ 50) kg | 8.1 g | / KCTI-CQI-20102
|  | (50 ~ 200) kg | 74 g
|  | (200 ~ 500) kg | 0.31 kg
Counter beam balances | 20105 | (0 ~ 310) g | 29 mg | Weight
|  | (310 ~ 2 610) g | 62 mg | / KCTI-CQI-20105
|  | (2 610 ~ 20 000) g | 8.6 g
Electric balances | 20109 | (0 ~ 200) g | 0.20 mg | Weight
|  | (200 ~ 500) g | 0.51 mg | / KCTI-CQI-20109
|  | (500 ~ 1 000) g | 0.53 mg
|  | (1 ~ 3) kg | 2.0 mg
|  | (3 ~ 5) kg | 2.8 mg
|  | (5 ~ 10) kg | 10 mg
|  | (10 ~ 30) kg | 77 mg
|  | (30 ~ 50) kg | 0.75 g
|  | (50 ~ 200) kg | 1.5 g
|  | (200 ~ 300) kg | 4.4 g
|  | (300 ~ 500) kg | 7.0 g
|  | (500 ~ 1 000) kg | 45 g
|  | (1 000 ~ 3 000) kg | 85 g
Platform scale balances | 20112 | (0 ~ 50) kg | 9.1 g | Weight
|  | (50 ~ 200) kg | 91 g | / KCTI-CQI-20112
|  | (200 ~ 500) kg | 0.19 kg
Spring scale balances | 20113 | (0 ~ 1) kg | 0.64 g
|  | (1 ~ 5) kg | 3.1 g | Weight
|  | (5 ~ 10) kg | 6.2 g | / KCTI-CQI-20113
|  | (10 ~ 50) kg | 58 g
|  | (50 ~ 100) kg | 0.31 kg
201. Mass
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Weights | 20116 | (1 mg ~ 10 kg) | F1 class | Weight,Electric balances
|  | 1 mg | 2.6 μg | / KCTI-CQI-20116
|  | 2 mg | 2.4 μg
|  | 5 mg | 2.4 μg
|  | 10 mg | 3.0 μg
|  | 20 mg | 3.5 μg
|  | 50 mg | 4.3 μg
|  | 100 mg | 5.3 μg
|  | 200 mg | 6.2 μg
|  | 500 mg | 8.2 μg
|  | 1 g | 9.1 μg
|  | 2 g | 12 μg
|  | 5 g | 15 μg
|  | 10 g | 18 μg
|  | 20 g | 24 μg
|  | 50 g | 0.21 mg
|  | 100 g | 0.21 mg
|  | 200 g | 0.16 mg
|  | 500 g | 0.30 mg
|  | 1 kg | 0.6 mg
|  | 2 kg | 2.3 mg
|  | 5 kg | 3.3 mg
|  | 10 kg | 6.4 mg
|  | (20 kg) | M1 class
|  | 20 kg | 0.15 g
202. Force
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Tension/Compression testing | 20203
machines
Tension |  | (20 ~ 200) N | 1.6 × 10-3 | Weight
|  | (0.2 ~ 5) kN | 1.2 × 10-3 | Force calibration machines
|  | (5 ~ 10) kN | 1.2 × 10-3 | / KCTI-CQI-20203
Compression |  | (20 ~ 500) N | 1.3 × 10-3
|  | (0.5 ~ 20) kN | 1.3 × 10-3
|  | (20 ~ 50) kN | 1.8 × 10-3
|  | (50 ~ 100) kN | 2.0 × 10-3
|  | (100 ~ 200) kN | 2.3 × 10-3
|  | (200 ~ 500) kN | 1.7 × 10-3
|  | (0.5 ~ 1) MN | 2.4 × 10-3
|  | (1 ~ 2) MN | 3.2 × 10-3
|  | (2 ~ 3) MN | 1.7 × 10-3
Push-pull gauges | 20204 | (0.4 ~ 500) N | 1.1 × 10-3 | Weight
|  | (500 ~ 1 000) N | 7.5 × 10-4 | / KCTI-CQI-20204
203. Torque
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Torque measuring devices | 20302 | (0.1 ~ 1) N·m | 2.8 × 10-3 | Torque arms,Weights
|  | (1 ~ 10) N·m | 2.6 × 10-3 | / KCTI-CQI-20302
|  | (10 ~ 100) N·m | 1.8 × 10-3
|  | (100 ~ 200) N·m | 1.4 × 10-3
|  | (200 ~ 500) N·m | 1.3 × 10-3
|  | (500 ~ 1 000) N·m | 1.7 × 10-3
|  | (1 000 ~ 1 500) N·m | 1.7 × 10-3
Torque wrenches/drivers | 20303 | (0.2 ~ 50) N·m | 1.0 × 10-2 | Torque calibration machine
|  | (50 ~ 200) N·m | 1.1 × 10-2 | / KCTI-CQI-20303
|  | (200 ~ 500) N·m | 1.0 × 10-2
204. Pressure
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Absolute pressure gauges | 20406
Barometer |  | (80 ~ 120) kPa | 2.8 × 10 -4 | Digital pressure gauge
/ KCTI-CQI-20406
Compound pressure gauge | 20408 | (-95 ~ 2 500) kPa | 7.5 × 10 -4 | Digital pressure gauge
/ KCTI-CQI-20408
Differential pressure gauge | 20409 | (0 ~ 0.2) kPa | 5.0 × 10 -3 | Digital pressure gauge
|  | (0.2 ~ 2.5) kPa | 7.0 × 10 -4 | Air Dead Weight Tester
|  | (2.5 ~ 600) kPa | 4.0 × 10 -3 | / KCTI-CQI-20409
Gauge pressure gauge ; | 20411
Dial,digital pressure gauge |  | (0 ~ 0.1) MPa | 2.0 × 10 -4 | Air/Oil Dead Weight Tester
recorders |  | (0.1 ~ 100) MPa | 1.0 × 10 -4 | Digital pressure gauge
/ KCTI-CQI-20411
Pressure transducers/ | 20412 | (0 ~ 0.1) MPa | 7.4 × 10 -4 | Air/Oil Dead Weight Tester
transmitters |  | (0.1 ~ 100) MPa | 7.2 × 10 -4 | Digital pressure gauge
Multi meter
/ KCTI-CQI-20412
Dial type vacuum gauge | 20413 | (-95 ~ 0) kPa | 2.2 × 10 -3 | Digital pressure gauge
/ KCTI-CQI-20413
206. Volume
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Volumetric glassware | 20601 | (0 ~ 1) ml | 0.63 μl | Electric balances
|  | (1 ~ 2) ml | 0.95 μl | / KCTI-CQI-20601
|  | (2 ~ 5) ml | 1.7 μl
|  | (5 ~ 10) ml | 2.3 μl
|  | (10 ~ 25) ml | 4.2 μl
|  | (25 ~ 50) ml | 7.0 μl
|  | (50 ~ 100) ml | 14 μl
|  | (100 ~ 200) ml | 29 μl
|  | (200 ~ 250) ml | 55 μl
|  | (250 ~ 500) ml | 86 μl
|  | (500 ~ 1 000) ml | 0.17 ml
|  | (1 000 ~ 2 000) ml | 0.33 ml
206. Volume
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Concrete air content meter | 20605 | (0 ~ 10 ) % | 0.05 % | Electric balances,Weight
/ KCTI-CQI-20605
Piston type volume meters | 20606 | (0.001 ~ 0.01) ml | 18 nl | Electric balances
|  | (0.01 ~ 0.02) ml | 40 nl | / KCTI-CQI-20606
|  | (0.02 ~ 0.05) ml | 70 nl
|  | (0.05 ~ 0.1) ml | 92 nl
|  | (0.1 ~ 0.2) ml | 0.20 μl
|  | (0.2 ~ 0.5) ml | 0.36 μl
|  | (0.5 ~ 1) ml | 0.66 μl
|  | (1 ~ 2) ml | 1.3 μl
|  | (2 ~ 5) ml | 3.1 μl
|  | (5 ~ 10) ml | 6.1 μl
|  | (10 ~ 20) ml | 13 μl
|  | (20 ~ 50) ml | 31 μl
207. Density
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Liquid density meters | 20702 | (0.75 ~ 1.5) g/cm3 | 0.000 10 g/cm3 | Density standard
/ KCTI-CQI-20702
Salinity meters | 20704 | (0 ~ 26) % | 0.042% | NaCl
/ KCTI-CQI-20704
Sucrose meters | 20705 | (0 ~ 60) % | 0.044% | Sucrose
/ KCTI-CQI-20705
Chloride meters | 20707 | (0 ~ 0.5) % | 0.0008% | Density standard
|  | (0.5 ~ 1.0) % | 0.0009% | / KCTI-CQI-20707
|  | (1.0 ~ 1.5) % | 0.0011%
208. Viscosity
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Kinematic viscometers; | 20801
capillary, etc
Zahn cup |  | (2.5 ~ 1 000) mm2/s | 2.6 × 10-2 | Viscosity standardsv
/ KCTI-CQI-20801
Ford cup |  | (2.5 ~ 1 000) mm2/s | 2.6 × 10-2
Dynamic viscometers; | 20802
rotaional, etc.
Rotational viscometer |  | (20 ~ 200 000) mPa·s | 1.6 × 10-2 | Viscosity standards
/ KCTI-CQI-20802
210. Hardness
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Brinell hardness testers | 21001 | (100 ~ 250) HBW 10/3 000 | 3.0 HBW 10 / 3 000 | Brinell Hardness test blocks
|  | (250 ~ 450) HBW 10/3 000 | 4.9 HBW 10 / 3 000 | Non-contact coordinate
measuring machines
/ KCTI-CQI-21001
Rockwell hardness testers | 21002 | (20 ~ 70) HRC | 0.46 HRC | Rockwell Hardness test
|  | (20 ~ 100) HRBW | 0.72 HRBW | blocks,Non-contact coordi-
nate measuring machines
/ KCTI-CQI-21002
Shore hardness testers | 21003 | (20 ~ 100) HS | 1.1 HS | Shore Hardness test blocks
/ KCTI-CQI-21003
Vickers hardness testers | 21004 | (40 ~ 240) HV 0.2 | 8.2 HV 0.2 | Vickers Hardness test blocks
|  | (240 ~ 600) HV 0.2 | 18 HV 0.2 | Non-contact coordinate
|  | (600 ~ 750) HV 0.2 | 39 HV 0.2 | measuring machines
/ KCTI-CQI-21004
210. Hardness
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Durometer hardness testers | 21005 | (0 ~ 100) HDA | 1.4 HDA | Durometer calibration
|  | (0 ~ 100) HDD | 1.1 HDD | / KCTI-CQI-21005
Leeb hardness testers | 21006 | 500 HLD 미만 | 4.6 | Hardness test blocks
|  | (500 ~ 700) HLD | 4.6 | / KCTI-CQI-21006
|  | 700 HLD 초과 | 4.6
301. Time/frequency
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Frequency meters/counters | 30104
Comparison with standard |  | 10 MHz | 7.2 × 10 -12 | GPS receiver,
frequency |  |  |  | Frequency counters
Reference oscillator |  | 10 MHz | 8.0 × 10 -13 | / KCTI-CQI-30104
Time interval meters/ | 30106
stop watches/timers
Stop watches & timers |  | 100 ms ~ 86 400 s | 4.2 × 10 -7 | Stopwatch calibrator
/ KCTI-CQI-30106
302. Velocity & revolution
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Standard RPM generators | 30201
speed of revolution |  | (6 ~ 60) min-1 | 1.2 min-1 | GPS Reciever, Universal Counter
|  | (60 ~ 600) min-1 | 0.13 min-1 | Standard Tachometer
|  | (600 ~ 6 000) min-1 | 0.44 min-1 | / KCTI-CQI-30201
|  | (6 000 ~ 60 000) min-1 | 4.3 min-1
|  | (60 000 ~ 90 000) min-1 | 6.4 min-1
Contact type tachometers | 30202 | (6 ~ 4 000) min-1 | 0.083 min-1 | GPS receiver,
RPM calibration system
/ KCTI-CQI-30202
Photo tachometers/stroboscopes | 30203
Photo tachometers/stroboscopes |  | (6 ~ 6 000) min-1 | 0.064 min-1 | GPS receiver, Optical
|  | (6 000 ~ 90 000) min-1 | 0.64 min-1 | rotation signal detector
/ KCTI-CQI-30203
flash generator |  | (6 ~ 6 000) min-1 | 0.022 min-1
|  | (6 000 ~ 90 000) min-1 | 0.63 min-1
401. DC volatage & current
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
DC ammeters | 40101
DC ammeters |  | 1 mA | 1.7 × 10 -4 | Meter Calibrator
|  | (1 ~ 10) mA | 1.5 × 10 -4 | / KCTI-CQI-40101
|  | (10 ~ 100) mA | 1.5 × 10 -4
|  | (0.1 ~ 1) A | 2.8 × 10 -4
|  | (1 ~ 10) A | 6.4 × 10 -4
401. DC volatage & current
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
DC voltage / | 40103
current calibrators
DC voltage |  | (0 ~ 100) mV | 1.1 × 10 -5 | Digital Multimeter
|  | (0.1 ~ 1) V | 6.3 × 10 -6 | Current Shunt
|  | (1 ~ 10) V | 5.3 × 10 -6 | / KCTI-CQI-40103
|  | (10 ~ 100) V | 7.7 × 10 -6
|  | (100 ~ 1 000) V | 8.0 × 10 -6
DC current |  | (0 ~ 100) μA | 9.6 × 10 -5
|  | (0.1 ~ 1) mA | 3.7 × 10 -5
|  | (1 ~ 10) mA | 3.5 × 10 -5
|  | (10 ~ 100) mA | 6.3 × 10 -5
|  | (0.1 ~ 1) A | 2.3 × 10 -4
|  | (1 ~ 10) A | 5.2 × 10 -4
|  | (10 ~ 20) A | 4.9 × 10 -4
|  | (20 ~ 100) A | 2.0 × 10 -4
Electrical temperature | 40104
calibrators(Without sensor)
Thermocouple(SOURCE) |  | (-8.825 ~ 76.373) mV | 1.1 μV | Meter Calibrator
Digital multimeters
RTD (SOURCE) |  | (17.136 ~ 100)Ω | 0.92 mΩ | / KCTI-CQI-40104
|  | (0.1 ~ 0.5)kΩ | 5.0 mΩ
|  | (0.5 ~ 1)kΩ | 9.2 mΩ
|  | (1 ~ 3.137 1)kΩ | 34 mΩ
Thermocouple(MEASURE) |  | (-8.825 ~ 76.373) mV | 3.8 μV
RTD (MEASURE) |  | (17.136 ~ 100)Ω | 13 mΩ
|  | (0.1 ~ 0.5)kΩ | 70 mΩ
|  | (0.5 ~ 1)kΩ | 82 mΩ
|  | (1 ~ 3.137 1)kΩ | 0.66 Ω
DC current shunts | 40105
Resistance |  | (0 ~ 1) mΩ | 1.5 × 10 -4 | Digital Multimeter
|  | (1 ~ 10) mΩ | 1.3 × 10 -3 | Meter Calibrator
|  | (10 ~ 100) mΩ | 1.2 × 10 -3 | Current Amplifier
|  | (0.1 ~ 1) Ω | 1.5 × 10 -4 | / KCTI-CQI-40105
|  | (1 ~ 10) Ω | 1.5 × 10 -4
|  | (10 ~ 100) Ω | 1.8 × 10 -4
|  | (0.1 ~ 1) kΩ | 7.0 × 10 -4
DC power supplies | 40108
DC voltage |  | (0 ~ 100) mV | 5.9 × 10 -5 | Meter Calibrator
|  | (0.1 ~ 1) V | 5.8 × 10 -5 | / KCTI-CQI-40108
|  | (1 ~ 10) V | 7.8 × 10 -6
|  | (10 ~ 100) V | 9.6 × 10 -6
|  | (100 ~ 1 000) V | 5.9 × 10 -5
DC current |  | (0 ~ 100) μA | 5.9 × 10 -4
|  | (0.1 ~ 1) mA | 5.8 × 10 -3
|  | (1 ~ 10) mA | 5.8 × 10 -4
|  | (10 ~ 100) mA | 8.6 × 10 -5
|  | (0.1 ~ 1) A | 1.2 × 10 -3
|  | (1 ~ 10) A | 1.5 × 10 -4
|  | (10 ~ 100) A | 2.0 × 10 -4
401. DC volatage & current
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
DC voltmeters | 40112
DC voltmeters |  | (1 ~ 100) mV | 3.8 × 10 -5 | Meter Calibrator
|  | (0.1 ~ 1) V | 1.8 × 10 -5 | / KCTI-CQI-40112
|  | (1 ~ 10) V | 1.8 × 10 -5
|  | (10 ~ 100) V | 2.4 × 10 -5
|  | (100 ~ 1 000) V | 2.6 × 10 -5
402. Resistance, capacitance and inductance
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Earth testers | 40205
Resistance |  | (0 ~ 100) mΩ | 4.1 × 10 -3 | High Voltage Meter
|  | (0.1 ~ 1) Ω | 7.1 × 10 -4 | Decade Resistance Box
|  | (0.001 ~ 100) kΩ | 7.1 × 10 -4 | Meter Calibrator
/ KCTI-CQI-40205
AC voltage |  | 50 ~ 100 Hz
|  | (1 ~ 100) V | 5.8 × 10 -3
|  | (100 ~ 300) V | 3.7 × 10 -4
Insulation testers | 40210
Insulation Resistance |  | (1 ~ 10) kΩ | 6.2 × 10 -4 | High Voltage Meter
|  | (10 ~ 100) kΩ | 6.2 × 10 -4 | Decade Resistance Box
|  | (0.1 ~ 1) MΩ | 6.2 × 10 -4 | Meter Calibrator
|  | (1 ~ 10) MΩ | 7.2 × 10 -4 | / KCTI-CQI-40210
|  | (10 ~ 100) MΩ | 7.2 × 10 -4
|  | (0.1 ~ 1) GΩ | 8.2 × 10 -3
|  | (1 ~ 10) GΩ | 1.3 × 10 -2
|  | (10 ~ 100) GΩ | 1.3 × 10 -2
|  | (0.1 ~ 1) TΩ | 1.3 × 10 -2
AC voltage |  | 60 Hz
|  | (10 ~ 100) V | 5.8 × 10 -3
|  | (100 ~ 200) V | 6.2 × 10 -4
|  | (200 ~ 300) V | 3.7 × 10 -4
|  | (300 ~ 400) V | 4.3 × 10 -4
|  | (400 ~ 500) V | 4.3 × 10 -4
|  | (500 ~ 600) V | 4.0 × 10 -4
|  | (600 ~ 1 000) V | 6.8 × 10 -4
DC output voltage |  | (1 ~ 10) V | 6.0 × 10 -5
|  | (10 ~ 50) V | 6.0 × 10 -5
|  | (50 ~ 100) V | 6.0 × 10 -5
|  | (100 ~ 500) V | 6.0 × 10 -5
|  | (0.5 ~ 1) kV | 6.2 × 10 -5
|  | (1 ~ 5) kV | 6.2 × 10 -3
402. Resistance, capacitance and inductance
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Resistance meter, high | 40214
resistance meter, etc.
resistance |  | 1 mΩ | 6.4 × 10 -5 | Decade Resistance Box
|  | (1 ~ 10) mΩ | 6.4 × 10 -5 | Standard Resistor
|  | (10 ~ 100) mΩ | 2.1 × 10 -5 | / KCTI-CQI-40214
|  | (0.1 ~ 1) Ω | 1.6 × 10 -5
|  | (1 ~ 10) Ω | 1.3 × 10 -5
|  | (10 ~ 100) Ω | 1.2 × 10 -5
|  | (0.1 ~ 1) kΩ | 1.1 × 10 -5
|  | (1 ~ 10) kΩ | 1.1 × 10 -5
|  | (10 ~ 100) kΩ | 1.4 × 10 -4
|  | (0.1 ~ 1) MΩ | 6.2 × 10 -4
|  | (1 ~ 10) MΩ | 7.2 × 10 -4
|  | (10 ~ 100) MΩ | 7.2 × 10 -4
Resistor, standard resistance, | 40215
stepped resistance, high
resistance, etc.
resistance |  | (0 ~ 10) mΩ | 1.1 × 10 -3 | Digital Multimeter
|  | (10 ~ 100) mΩ | 5.0 × 10 -4 | / KCTI-CQI-40215
|  | (0.1 ~ 1) Ω | 1.6 × 10 -4
|  | (1 ~ 10) Ω | 1.8 × 10 -5
|  | (10 ~ 100) Ω | 1.1 × 10 -5
|  | (0.1 ~ 100) kΩ | 1.1 × 10 -5
|  | (0.1 ~ 1) MΩ | 1.3 × 10 -5
|  | (1 ~ 10) MΩ | 5.0 × 10 -5
|  | (10 ~ 100) MΩ | 4.3 × 10 -4
|  | (0.1 ~ 1) GΩ | 4.7 × 10 -3
Impedance bridges | 40217
/LCR meters
1 kHz
Resistance |  | (0 ~ 1) Ω | 7.3 × 10 -4 | Standard Capacitor
|  | (1 ~ 10) Ω | 6.8 × 10 -4 | Standard Inductor
|  | (10 ~ 100) Ω | 6.5 × 10 -4 | / KCTI-CQI-40217
|  | (0.1 ~ 1) kΩ | 2.9 × 10 -4
|  | (1 ~ 100) kΩ | 1.2 × 10 -4
|  | (0.1 ~ 1) MΩ | 2.9 × 10 -4
Capacitance |  | 1 kHz
|  | (0 ~ 1 000) pF | 3.5 × 10 -4
|  | (1 ~ 10) nF | 7.6 × 10 -5
|  | (0.01 ~ 1) μF | 1.1 × 10 -4
Inductance |  | 1 kHz
|  | (0 ~ 100) μH | 1.2 × 10 -3
|  | 100 μH ~ 1 H | 3.1 × 10 -4
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC ammeters | 40301
AC current |  | 50 Hz ~ 1 kHz |  | Meter Calibrator
|  | (1 ~ 100) mA | 1.4 × 10 -3 | / KCTI-CQI-40301
|  | (0.1 ~ 1) A | 8.5 × 10 -4
|  | (1 ~ 10) A | 2.4 × 10 -3
Clamp ammeters/voltmeters | 40302
DC voltage |  | (10 ~ 100) mV | 1.6 × 10 -4 | Current Coil
|  | (0.1 ~ 1) V | 3.8 × 10 -5 | Meter Calibrator
|  | (1 ~ 1 000) V | 2.4 × 10 -5 | / KCTI-CQI-40302
AC voltage |  | 40 Hz
|  | (10 ~ 100) mV | 1.8 × 10 -3
|  | (0.1 ~ 1) V | 4.5 × 10 -4
|  | (1 ~ 10) V | 4.3 × 10 -4
Current Coil
|  | 50 Hz ~ 1 kHz |  | Meter Calibrator
|  | (10 ~ 100) mV | 1.1 × 10 -3 | / KCTI-CQI-40302
|  | (0.1 ~ 1) V | 2.8 × 10 -4
|  | (1 ~ 1 000) V | 3.7 × 10 -4
DC current |  | (1 ~ 100) mA | 1.9 × 10 -4
|  | (0.1 ~ 10) A | 1.3 × 10 -3
|  | (10 ~ 1 000) A | 3.0 × 10 -3
AC current |  | 50 Hz ~ 1 kHz
|  | (1 ~ 100) mA | 1.4 × 10 -3
|  | (0.1 ~ 1) A | 8.5 × 10 -4
|  | (1 ~ 10) A | 2.4 × 10 -3
50 Hz ~ 100 Hz
|  | (1 ~ 100) mA | 1.4 × 10 -3
|  | (0.1 ~ 1) A | 8.5 × 10 -4
|  | (1 ~ 10) A | 2.0 × 10 -3
|  | (10 ~ 1 000) A | 5.4 × 10 -3
Resistance |  | (1 ~ 100) Ω | 1.2 × 10 -3
|  | (0.1 ~ 100) kΩ | 1.3 × 10 -4
|  | (0.1 ~ 10) MΩ | 1.8 × 10 -4
|  | (10 ~ 100) MΩ | 6.4 × 10 -4
AC voltage/current | 40303
calibrators
AC voltage |  | 40 Hz ~ 100 Hz |  | Digital Multimeter
|  | (10 ~ 100) mV | 2.6 × 10 -4 | Current Shunt
|  | (0.1 ~ 1) V | 1.6 × 10 -4 | Frequency Counter
|  | (1 ~ 10) V | 1.6 × 10 -4 | / KCTI-CQI-40303
|  | (10 ~ 100) V | 1.6 × 10 -4
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC voltage/current | 40303 | 50 Hz ~ 100 Hz |  | Digital Multimeter
calibrators |  | (100 ~ 1 000) V | 1.5 × 10 -4 | Current Shunt
Frequency Counter
|  | 100 Hz ~ 1 kHz |  | / KCTI-CQI-40303
|  | (10 ~ 100) mV | 2.2 × 10 -4
|  | (0.1 ~ 1) V | 1.2 × 10 -4
|  | (1 ~ 10) V | 1.2 × 10 -4
|  | (10 ~ 100) V | 1.2 × 10 -4
|  | (100 ~ 1 000) V | 1.5 × 10 -4
AC current |  | 40 Hz ~ 100 Hz
|  | (0.1 ~ 1) mA | 5.6 × 10 -4
|  | (1 ~ 100) mA | 5.6 × 10 -4
|  | (0.1 ~ 1) A | 1.1× 10 -3
|  | (1 ~ 10) A | 1.3× 10 -3
|  | (10 ~ 100) A | 3.3× 10 -3
50 Hz ~ 100 Hz
|  | (10 ~ 100) A | 3.3× 10 -3
100 Hz ~ 1 kHz
|  | (0.1 ~ 1) mA | 5.6 × 10 -4
|  | (1 ~ 100) mA | 5.6 × 10 -4
|  | (0.1 ~ 1) A | 1.1× 10 -3
|  | (1 ~ 10) A | 1.4× 10 -3
|  | (10 ~ 100) A | 3.3× 10 -3
AC current shunts | 40305
AC Resistance |  | 60 Hz ~ 1 kHz |  | Digital Multimeter
|  | (1 ~ 10) mΩ | 2.4 × 10 -3 | Current Shunt
|  | (10 ~ 100) mΩ | 8.8 × 10 -4 | Frequency Counter
|  | (0.1 ~ 1) Ω | 7.6 × 10 -4 | / KCTI-CQI-40305
|  | (1 ~ 10) Ω | 7.6 × 10 -4
|  | (10 ~ 100) Ω | 1.4 × 10 -3
|  | (0.1 ~ 1) kΩ | 3.0 × 10 -3
Power factor meters | 40310
Power factor |  | 50 Hz, 60 Hz |  | Meter Calibrator
(LEAD & LAG) |  | 1 | 0.001 | / KCTI-CQI-40310
|  | 0.9 | 0.002
|  | 0.8 | 0.002
|  | 0.7 | 0.003
|  | 0.6 | 0.003
|  | 0.5 | 0.004
|  | 0.4 | 0.005
|  | 0.3 | 0.007
|  | 0.2 | 0.01
|  | 0.1 | 0.021
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC power meters | 40311
AC voltage |  | 50 Hz, 60 Hz |  | Meter Calibrator
|  | 10 mV | 1.0 × 10 -3 | / KCTI-CQI-40311
|  | (10 ~ 100) mV | 2.7 × 10 -4
|  | (0.1 ~ 1) V | 2.5 × 10 -4
|  | (1 ~ 10) V | 2.5 × 10 -4
|  | (10 ~ 100) V | 2.5 × 10 -4
|  | (100 ~ 1 000) V | 3.6 × 10 -4
AC current |  | 50 Hz, 60 Hz
|  | 100 μA | 2.7 × 10 -3
|  | (0.1 ~ 1) mA | 1.4 × 10 -3
|  | (1 ~ 10) mA | 7.0 × 10 -4
|  | (10 ~ 100) mA | 7.0 × 10 -4
|  | (0.1 ~ 1) A | 7.0 × 10 -4
|  | (1 ~ 10) A | 9.3 × 10 -4
|  | (10 ~ 20) A | 1.7 × 10 -3
Power |  | (240V, 50 Hz)
|  | (0.24 ~ 1.2) W | 2.1 × 10 -2
|  | (1.2 ~ 2.4) W | 2.1 × 10 -2
|  | (2.4 ~ 6) W | 1.0 × 10 -2
|  | (6 ~ 12) W | 2.1 × 10 -2
|  | (12 ~ 24) W | 2.1 × 10 -2
|  | (24 ~ 60) W | 1.0 × 10 -2
|  | (60 ~ 120) W | 2.5 × 10 -2
|  | (120 ~ 240) W | 2.1 × 10 -2
|  | (240 ~ 600) W | 2.1 × 10 -2
|  | (0.6 ~ 1.2) kW | 1.0 × 10 -2
|  | (1.2 ~ 2.4) kW | 5.2 × 10 -3
|  | (2.4 ~ 4.8) kW | 2.7 × 10 -3
(240V, 60 Hz)
|  | (0.24 ~ 1.2) W | 2.1 × 10 -2
|  | (1.2 ~ 2.4) W | 2.1 × 10 -2
|  | (2.4 ~ 6) W | 1.0 × 10 -2
|  | (6 ~ 12) W | 2.1 × 10 -2
|  | (12 ~ 24) W | 2.1 × 10 -2
|  | (24 ~ 60) W | 1.0 × 10 -2
|  | (60 ~ 120) W | 2.5 × 10 -2
|  | (120 ~ 240) W | 2.1 × 10 -2
|  | (240 ~ 600) W | 2.1 × 10 -2
|  | (0.6 ~ 1.2) kW | 1.0 × 10 -2
|  | (1.2 ~ 2.4) kW | 5.2 × 10 -3
|  | (2.4 ~ 4.8) kW | 3.1 × 10 -3
(120V, 50 Hz)
|  | (0.12 ~ 0.6) W | 2.2 × 10 -2
|  | (0.6 ~ 1.2) W | 2.5 × 10 -2
|  | (1.2 ~ 3) W | 1.3 × 10 -2
|  | (3 ~ 6) W | 2.2 × 10 -2
|  | (6 ~ 12) W | 2.5 × 10 -2
|  | (12 ~ 30) W | 1.3 × 10 -2
|  | (30 ~ 60) W | 2.2 × 10 -2
|  | (60 ~ 120) W | 2.5 × 10 -2
|  | (120 ~ 300) W | 2.1 × 10 -2
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC power meters | 40311
Power |  | (300 ~ 600) W | 1.0 × 10 -2 | Meter Calibrator
|  | (0.6 ~ 1.2) kW | 4.2 × 10 -3 | / KCTI-CQI-40310
|  | (1.2 ~ 2.4) kW | 3.0 × 10 -3
(120V, 60 Hz)
|  | (0.12 ~ 0.6) W | 2.2 × 10 -2
|  | (0.6 ~ 1.2) W | 2.5 × 10 -2
|  | (1.2 ~ 3) W | 1.3 × 10 -2
|  | (3 ~ 6) W | 2.2 × 10 -2
|  | (6 ~ 12) W | 2.5 × 10 -2
|  | (12 ~ 30) W | 1.3 × 10 -2
|  | (30 ~ 60) W | 2.2 × 10 -2
|  | (60 ~ 120) W | 2.5 × 10 -2
|  | (120 ~ 300) W | 2.1 × 10 -2
|  | (300 ~ 600) W | 1.0 × 10 -2
|  | (0.6 ~ 1.2) kW | 4.2 × 10 -3
|  | (1.2 ~ 2.4) kW | 3.0 × 10 -3
Power factor |  | 50 Hz, 60 Hz
(LEAD & LAG) |  | 1 | 0.001
|  | 0.9 | 0.002
|  | 0.8 | 0.002
|  | 0.7 | 0.003
|  | 0.6 | 0.003
|  | 0.5 | 0.004
|  | 0.4 | 0.005
|  | 0.3 | 0.007
|  | 0.2 | 0.01
|  | 0.1 | 0.021
AC power supplies | 40312 |  |  | Digital Multimeter
AC voltage |  | 40 Hz ~ 100 Hz |  | Frequency Counter
|  | 1 V | 0.058 V | AC Load
|  | (1 ~ 10) V | 5.8 × 10 -3 | Current Shunt
|  | (10 ~ 100) V | 6.0 × 10 -4 | / KCTI-CQI-40312
|  | (100 V ~ 1 000) V | 1.6 × 10 -4
100 Hz ~ 1 kHz
|  | 1 V | 0.058 V
|  | (1 ~ 10) V | 5.8 × 10 -3
|  | (10 ~ 100) V | 5.9 × 10 -4
|  | (100 V ~ 1 000) V | 1.6 × 10 -4
AC current |  | 40 Hz ~ 1 kHz
|  | 100 mA | 0.14 mA
|  | (0.1 ~ 1) A | 1.4 × 10 -3
|  | (1 ~ 10) A | 3.2 × 10 -3
|  | (10 ~ 20) A | 6.2 × 10 -3
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Puncture/safety testers | 40313
DC voltage |  | (0 ~ 0.5) kV | 1.0 × 10 -2 | Decade Resistance Box
|  | (0.5 ~ 1) kV | 2.0 × 10 -3 | Digital Multimeter
|  | (1 ~ 5) kV | 7.0 × 10 -3 | High Voltage Meter
|  | (5 ~ 10) kV | 6.4 × 10 -3 | / KCTI-CQI-40313
|  | (10 ~ 20) kV | 1.8 × 10 -2
|  | (20 ~ 50) kV | 2.0 × 10 -2
|  | (50 ~ 100) kV | 1.9 × 10 -2
AC voltage |  | 50 Hz
|  | (0.1 ~ 0.5) kV | 1.0 × 10 -2
|  | (0.5 ~ 1) kV | 2.0 × 10 -3
|  | (1 ~ 5) kV | 7.0 × 10 -3
|  | (5 ~ 10) kV | 6.4 × 10 -3
60 Hz
|  | (0.1 ~ 0.5) kV | 1.0 × 10 -2
|  | (0.5 ~ 1) kV | 2.0 × 10 -3
|  | (1 ~ 5) kV | 7.0 × 10 -3
|  | (5 ~ 10) kV | 6.4 × 10 -3
|  | (10 ~ 20) kV | 1.9 × 10 -2
|  | (20 ~ 50) kV | 2.2 × 10 -2
|  | (50 ~ 100) kV | 2.0 × 10 -2
DC current |  | (0 ~ 0.5) mA | 2.0 × 10 -3
|  | (0.5 ~ 1) mA | 1.0 × 10 -3
|  | (1 ~ 2) mA | 5.0 × 10 -4
|  | (2 ~ 5) mA | 2.0 × 10 -4
|  | (5 ~ 10) mA | 1.0 × 10 -4
|  | (10 ~ 20) mA | 2.5 × 10 -4
|  | (20 ~ 50) mA | 1.2 × 10 -4
|  | (50 ~ 100) mA | 7.0 × 10 -5
AC current |  | 50 Hz
|  | (0 ~ 0.5) mA | 1.5 × 10 -3
|  | (0.5 ~ 1) mA | 1.0 × 10 -3
|  | (1 ~ 2) mA | 2.0 × 10 -3
|  | (2 ~ 5) mA | 1.0 × 10 -3
|  | (5 ~ 10) mA | 6.0 × 10 -4
|  | (10 ~ 20) mA | 1.8 × 10 -3
|  | (20 ~ 50) mA | 8.6 × 10 -4
|  | (50 ~ 100) mA | 5.6 × 10 -4
60 Hz
|  | (0 ~ 0.5) mA | 2.0 × 10 -3
|  | (0.5 ~ 1) mA | 1.0 × 10 -3
|  | (1 ~ 2) mA | 2.0 × 10 -3
|  | (2 ~ 5) mA | 1.0 × 10 -3
|  | (5 ~ 10) mA | 6.0 × 10 -4
|  | (10 ~ 20) mA | 1.8 × 10 -3
|  | (20 ~ 50) mA | 8.6 × 10 -4
|  | (50 ~ 100) mA | 5.6 × 10 -4
403. AC voltage, current & power
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
AC voltmeter/potential | 40318
difference, effective value
AC voltage |  | 40 Hz |  | Meter Calibrator
|  | (10 ~ 100) mV | 1.8 × 10 -3 | / KCTI-CQI-40318
|  | (0.1 ~ 1) V | 4.6 × 10 -4
50 Hz ~ 1 kHz
|  | (1 ~ 10) mV | 4.4 × 10 -4
|  | (10 ~ 100) mV | 1.1 × 10 -3
|  | (0.1 ~ 1) V | 2.8 × 10 -4
|  | (1 ~ 100) V | 2.6 × 10 -4
|  | (100 ~ 1 000) V | 3.7 × 10 -4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Line frequency meters | 40410
Frequency |  | 45 Hz | 2.6 × 10 -4 | Meter Calibrator
|  | (0.045 ~ 10) kHz | 5.9 × 10 -4 | / KCTI-CQI-40410
Leakage current testers | 40416
DC current |  | 10 μA | 2.5 × 10 -3 | Meter Calibrator
|  | (10 ~ 100) μA | 7.1 × 10 -4 | / KCTI-CQI-40416
|  | (0.1 ~ 100) mA | 6.1 × 10 -4
AC current |  | 50 Hz ~ 1 kHz
|  | 10 μA | 1.2 × 10 -2
|  | (10 ~ 100) μA | 1.4 × 10 -3
|  | (0.1 ~ 100) mA | 6.3 × 10 -4
AC voltmeters |  | 50 Hz ~ 1 kHz
|  | 10 mV | 1.1 × 10 -3 | Meter Calibrator
|  | 100 mV | 6.4 × 10 -4 | / KCTI-CQI-40416
|  | (0.1 ~ 100) V | 6.3 × 10 -4
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Electronic AC/DC | 40417
DC voltmeters |  | 100 mV | 6.7 × 10 -5 | Meter Calibrator
|  | (0.1 ~ 1 000) V | 6.2 × 10 -5 | / KCTI-CQI-40417
DC current |  | 10 μA | 2.5 × 10 -3
|  | (0.01 ~ 100) mA | 4.1 × 10 -4
|  | (0.1 ~ 1) A | 2.8 × 10 -4
|  | (1 ~ 10) A | 6.4 × 10 -4
|  | (10 ~ 100) A | 1.6 × 10 -4
(50 ~ 400 )Hz
AC voltmeters |  | 100 mV | 2.8 × 10 -4
|  | (0.1 ~ 100) V | 2.5 × 10 -4
|  | (100 ~ 1 000) V | 4.0 × 10 -4
(50 ~ 400 )Hz
AC current |  | 10 μA | 1.7 × 10 -2
|  | (0.01 ~ 100) mA | 1.8 × 10 -3
|  | (0.1 ~ 1) A | 1.4 × 10 -4
|  | (1 ~ 10) A | 2.5 × 10 -4
Analogue/digital multimeters ; | 40419
Individual attribute recognition
DC voltmeters |  | (1 ~ 100) mV | 3.7 × 10 -5 | Meter Calibrator
|  | (0.1 ~ 1) V | 1.6 × 10 -5 | / KCTI-CQI-40419
|  | (1 ~ 10) V | 1.7 × 10 -5
|  | (10 ~ 100) V | 2.3 × 10 -5
|  | (100 ~ 1 000) V | 2.3 × 10 -5
AC voltmeters |  | 40 Hz ~ 1 kHz
|  | (1 ~ 10) mV | 1.6 × 10 -3
|  | (10 ~ 100) mV | 4.5 × 10 -4
|  | (0.1 ~ 1) V | 4.1 × 10 -4
|  | (1 ~ 10) V | 4.3 × 10 -4
AC voltmeters |  | 1 kHz ~ 20 kHz
|  | (1 ~ 10) mV | 9.4 × 10 -4
|  | (10 ~ 100) mV | 2.9 × 10 -4
|  | (0.1 ~ 1) V | 2.9 × 10 -4
|  | (1 ~ 10) V | 3.5 × 10 -4
20 kHz ~ 100 kHz
|  | (1 ~ 10) mV | 5.4 × 10 -3
|  | (10 ~ 100) mV | 1.3 × 10 -3
|  | (0.1 ~ 1) V | 9.7 × 10 -4
|  | (1 ~ 10) V | 1.2 × 10 -3
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Analogue/digital multimeters ; | 40419 | 50 Hz ~ 1 kHz
Individual attribute recognition |  | (1 ~ 100) V | 2.5 × 10 -4 | Meter Calibrator
|  | (100 ~ 1 000) V | 3.6 × 10 -4 | / KCTI-CQI-40419
1 kHz ~ 20 kHz
|  | (10 ~ 100) V | 3.8 × 10 -4
DC current |  | (0.1 ~ 1) mA | 1.7 × 10 -4
|  | (1 ~ 10) mA | 1.4 × 10 -4
|  | (10 ~ 100) mA | 1.4 × 10 -4
|  | (0.1 ~ 1) A | 2.8 × 10 -4
|  | (1 ~ 10) A | 6.4 × 10 -4
AC current |  | 100 Hz ~ 1 kHz
|  | (0.1 ~ 1) mA | 2.3 × 10 -4
|  | (1 ~ 10) mA | 2.8 × 10 -4
|  | (10 ~ 100) mA | 2.9 × 10 -4
|  | (0.1 ~ 1) A | 2.6 × 10 -4
50 Hz ~ 1 kHz
|  | (1 ~ 10) A | 3.9 × 10 -4
Resistance |  | (1 ~ 10) Ω | 1.1 × 10 -4
|  | (10 ~ 100) Ω | 1.1 × 10 -4
|  | (0.1 ~ 1) kΩ | 5.1 × 10 -5
|  | (1 ~ 10) kΩ | 5.1 × 10 -5
|  | (10 ~ 100) kΩ | 6.0 × 10 -5
|  | (0.1 ~ 1) MΩ | 5.4 × 10 -5
|  | (1 ~ 10) MΩ | 1.6 × 10 -4
|  | (10 ~ 100) MΩ | 6.4 × 10 -4
Oscilloscopes | 40421
DC voltmeters |  | 1 mV | 4.7 × 10 -2 | Oscilloscope Calibrator
|  | (1 ~ 2) mV | 2.4 × 10 -2 | Frequency Counter
|  | (2 ~ 5) mV | 1.0 × 10 -2 | Digital Multimeter
|  | (5 ~ 10) mV | 5.2 × 10 -3 | Oscilloscope
|  | (10 ~ 20) mV | 2.9 × 10 -3 | / KCTI-CQI-40421
|  | (20 ~ 50) mV | 1.6 × 10 -3
|  | (50 ~ 100) mV | 1.1 × 10 -3
|  | (0.1 ~ 0.2) V | 1.0 × 10 -3
|  | (0.2 ~ 0.5) V | 8.0 × 10 -4
|  | (0.5 ~ 1) V | 7.0 × 10 -4
|  | (1 ~ 2) V | 6.5 × 10 -4
|  | (2 ~ 5) V | 6.0 × 10 -4
|  | (5 ~ 10) V | 6.0 × 10 -4
|  | (10 ~ 20) V | 1.0 × 10 -3
|  | (20 ~ 50) V | 8.0 × 10 -4
|  | (50 ~ 100) V | 6.0 × 10 -4
AC voltmeters |  | 2 mV | 2.5 × 10 -2
|  | (2 ~ 5) mV | 1.1 × 10 -2
|  | (5 ~ 10) mV | 5.8 × 10 -3
|  | (10 ~ 20) mV | 3.5 × 10 -3
|  | (20 ~ 50) mV | 2.2 × 10 -3
|  | (50 ~ 100) mV | 1.7 × 10 -3
|  | (0.1 ~ 0.2) V | 1.5 × 10 -3
|  | (0.2 ~ 0.5) V | 1.4 × 10 -3
|  | (0.5 ~ 1) V | 1.3 × 10 -3
|  | (1 ~ 2) V | 1.2 × 10 -3
|  | (2 ~ 5) V | 1.2 × 10 -3
|  | (5 ~ 10) V | 1.2 × 10 -3
|  | (10 ~ 20) V | 1.5 × 10 -3
|  | (20 ~ 50) V | 1.2 × 10 -3
|  | (50 ~ 100) V | 1.2 × 10 -3
404. Other DC & LF measurements
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Oscilloscopes | 40421
Time Marker |  | 5 ns | 2.0 × 10 -4 | Oscilloscope Calibrator
|  | (5 ~ 10) ns | 1.0 × 10 -3 | Frequency Counter
|  | (10 ~ 20) ns | 5.0 × 10 -4 | Digital Multimeter
|  | (20 ~ 50) ns | 2.0 × 10 -4 | Oscilloscope
|  | (50 ~ 100) ns | 1.0 × 10 -3 | / KCTI-CQI-40421
|  | (100 ~ 200) ns | 5.0 × 10 -4
|  | (200 ~ 500) ns | 2.0 × 10 -4
|  | (0.5 ~ 1) μs | 1.0 × 10 -3
|  | (1 ~ 2) μs | 5.0 × 10 -4
|  | (2 ~ 5) μs | 2.0 × 10 -4
|  | (5 ~ 10) μs | 1.0 × 10 -3
|  | (10 ~ 20) μs | 5.0 × 10 -4
|  | (20 ~ 50) μs | 2.0 × 10 -4
|  | (50 ~ 100) μs | 1.0 × 10 -3
|  | (100 ~ 200) μs | 5.0 × 10 -4
|  | (200 ~ 500) μs | 2.0 × 10 -4
|  | (0.5 ~ 1) ms | 1.0 × 10 -3
|  | (1 ~ 2) ms | 5.0 × 10 -4
|  | (2 ~ 5) ms | 2.0 × 10 -4
|  | (5 ~ 10) ms | 1.0 × 10 -3
|  | (10 ~ 20) ms | 5.0 × 10 -4
|  | (20 ~ 50) ms | 2.0 × 10 -4
|  | (50 ~ 100) ms | 1.0 × 10 -3
|  | (100 ~ 200) ms | 5.0 × 10 -4
|  | (200 ~ 500) ms | 8.0 × 10 -4
|  | (0.5 ~ 1) s | 2.0 × 10 -3
|  | (1 ~ 2) s | 2.5 × 10 -3
|  | (2 ~ 5) s | 6.0 × 10 -3
Level Sine Wave |  | 50 kHz | 1.7 × 10 -2
Ref. 50 kHz |  | (50 ~ 100) kHz | 1.7 × 10 -2 | Oscilloscope Calibrator
|  | (0.1 ~ 1) MHz | 1.7 × 10 -2 | Frequency Counter
|  | (1 ~ 10) MHz | 2.9 × 10 -2 | Digital Multimeter
|  | (10 ~ 50) MHz | 2.9 × 10 -2 | Oscilloscope
|  | (50 ~ 100) MHz | 2.9 × 10 -2 | / KCTI-CQI-40421
|  | (100 ~ 300) MHz | 2.9 × 10 -2
|  | (300 ~ 500) MHz | 2.9 × 10 -2
|  | (500 ~ 600) MHz | 3.7 × 10 -2
|  | (600 ~ 1 000) MHz | 3.7 × 10 -2
|  | (1 000 ~ 1 300) MHz | 4.7 × 10 -2
|  | (1 300 ~ 1 500) MHz | 4.7 × 10 -2
|  | (1 500 ~ 1 800) MHz | 4.7 × 10 -2
|  | (1 800 ~ 2 000) MHz | 4.7 × 10 -2
|  | (2 000 ~ 2 100) MHz | 4.7 × 10 -2
501. Contact thermometry
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Temperature generators; | 50101
ovens, furnaces, isothermal
liquid baths,ice-point baths,
dry-block calibrators
oven |  | (-40 ~ 200) ℃ | 0.9 ℃ | Standard thermometer, SPRT
temperature bridge
furnaces |  | (200 ~ 1 100) ℃ | 1.0 ℃ | temperature calibrator
/ KCTI-CQI-50101
isothermal liquid baths |  | (-80 ~ 200) ℃ | 0.06 ℃
dry-block calibrators |  | (-40 ~ 200) ℃ | 0.07 ℃
|  | (200 ~ 1 100) ℃ | 1.0 ℃
Temperature indicators | 50102
/recorders/
controllers,temperature
calibrators
Resistance(sensors included) |  | (-80 ~ -39) ℃ | 0.07 ℃ | Standard thermometer, SPRT
|  | (-39 ~ 200) ℃ | 0.08 ℃ | temperature bridge
temperature calibrator
Resistance |  | (-80 ~ -39) ℃ | 0.15 ℃ | / KCTI-CQI-50102
|  | (-39 ~ 200) ℃ | 0.3 ℃
Thermoelectric (including |  | (-80 ~ -39) ℃ | 0.15 ℃
sensor) |  | (-39 ~ 200) ℃ | 0.2 ℃
|  | (200 ~ 1 100) ℃ | 1.3 ℃
Thermoelectric
R-Type |  | (0 ~ 1 100) ℃ | 1.1 ℃ | Standard thermometer, SPRT
S-Type |  | (0 ~ 1 100) ℃ | 1.1 ℃ | temperature bridge
B-Type |  | (600 ~ 1 100) ℃ | 1.1 | temperature calibrator
K-Type |  | (-80 ~ -39) ℃ | 0.5 ℃ | / KCTI-CQI-50102
|  | (-39 ~ 1 100) ℃ | 0.5
N-Type |  | (0 ~ 1 100) ℃ | 0.6
E-Type |  | (-80 ~ -39) ℃ | 0.6 ℃
|  | (-39 ~ 700) ℃ | 0.6
J-Type |  | (-80 ~ -39) ℃ | 0.4 ℃
|  | (-39 ~ 600) ℃ | 0.4
T-Type |  | (-80 ~ -39) ℃ | 0.5 ℃
|  | (-39 ~ 300) ℃ | 0.5
Glass thermometers; | 50103
liquid-in-glass,Beckmann
liquid-in-glass |  | (-39 ~ 200) ℃ | 0.04 ℃ | temperature bridge,SPRT
/ KCTI-CQI-50103
Resistance thermometers; | 50104
SPRT, IPRT, thermistors, ect.
Industrial resistance
thermometers |  | (-80 ~ 200) ℃ | 0.06 ℃ | temperature bridge,SPRT
/ KCTI-CQI-50104
Thermal expansion thermomete- | 50105
rs;bimetal,gas or liquid type
bimetal |  | (-39 ~ 200) ℃ | 0.3 ℃ | temperature bridge,SPRT
/ KCTI-CQI-50105
501. Contact thermometry
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Thermomecouples; | 50106
noble metal, base metal, pure
metal, special type, ect.
base metal |  | (-80 ~ -39) ℃ | 0.06 ℃ | Standard thermomecouples
|  | (-39 ~ 200) ℃ | 0.3 ℃ | / KCTI-CQI-50106
|  | (200 ~ 1 100) ℃ | 1.3 ℃
Noble metal |  | (0 ~ 200) ℃ | 0.3 ℃ | Standard thermomecouples
|  | (200 ~ 1 100) ℃ | 1.6 ℃ | / KCTI-CQI-50106
Temperature transducers | 50107 | (-39 ~ 200) ℃ | 0.3 ℃ | Standard thermometer, SPRT
|  | (200 ~ 1 100) ℃ | 1.2 ℃ | temperature bridge
furnaces
/ KCTI-CQI-50107
502. Non contact thermometry
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Non contact thermometry | 50204
Standard radiation |  | (0 ~ 50) ℃ | 0.6 ℃ | Standard radiation
thermometers |  | (50 ~ 200) ℃ | 0.8 ℃ | thermometers
|  | (200 ~ 500) ℃ | 1.0 ℃ | / KCTI-CQI-50204
|  | (500 ~ 800) ℃ | 4.5 ℃
|  | (800 ~ 1 000) ℃ | 4.9 ℃
503. humidity
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Relative humidity hygrometers; | 50302
polimer thin film, hair, etc.
polimer thinfilm
humidity |  | (20 ~ 40) % R.H. | 2.1 % R.H. | Dew-point,TEMPERATURE &
|  | (40 ~ 60) % R.H. | 2.9 % R.H. | HUMIDITY CHAMBER
|  | (60 ~ 80) % R.H. | 3.2 % R.H. | / KCTI-CQI-50302
|  | (80 ~ 90) % R.H. | 3.8 % R.H.
temperature |  | (-40 ~ 20) ℃ | 1.1 ℃
|  | (20 ~ 90) ℃ | 1.4 ℃
hair
humidity |  | (20 ~ 90) % R.H. | 4.4 % R.H. | Dew-point,TEMPERATURE &
temperature |  | (10 ~ 90) ℃ | 1.9 ℃ | HUMIDITY CHAMBER
/ KCTI-CQI-50302
Psychrometers; | 50303
Assmann ventilated, PRT type,
ect.
PRT type, ect.
humidity |  | (20 ~ 90) % R.H. | 3.9 % R.H. | Dew-point,TEMPERATURE &
temperature |  | (10 ~ 90) ℃ | 1.2 ℃ | HUMIDITY CHAMBER
/ KCTI-CQI-50303
503. humidity
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Temperature humidity | 50304
recorders;
Hygrothermograph, ect.
humidity |  | (20 ~ 90) % R.H. | 3.7 % R.H. | Dew-point,TEMPERATURE &
temperature |  | (-20 ~ 50) ℃ | 2.3 ℃ | HUMIDITY CHAMBER
/ KCTI-CQI-50304
Transducers; | 50305
dew-point/relative humidity
humidity |  | (20 ~ 40) % R.H. | 2.5 % R.H. | Dew-point,TEMPERATURE &
|  | (40 ~ 60) % R.H. | 2.7 % R.H. | HUMIDITY CHAMBER
|  | (60 ~ 80) % R.H. | 3.4 % R.H. | / KCTI-CQI-50305
|  | (80 ~ 90) % R.H. | 4.0 % R.H.
Humidity generators; | 50306
two-pressure,two-temperature,
flow mixing humidity genera
tor,constant temperature and
humidity chamber, ect.
humidity |  | (20 ~ 40) % R.H. | 2.5 % R.H. | Dew-point,PRT SENSOR
|  | (40 ~ 60) % R.H. | 3.5 % R.H. | / KCTI-CQI-50306
|  | (60 ~ 80) % R.H. | 4.4 % R.H.
|  | (80 ~ 95) % R.H. | 5.0 % R.H.
temperature |  | (-40 ~ 0) ℃ | 0.8 ℃
|  | (0 ~ 90) ℃ | 1.0 ℃
601. Sound in air
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Sound level meters | 60106 | 125 Hz | 0.3 dB | Sound calibrators
|  | 250 Hz | 0.2 dB | / KCTI-CQI-60106
|  | 500 Hz | 0.2 dB
|  | 1 kHz | 0.2 dB
|  | 2 kHz | 0.2 dB
|  | 4 kHz | 0.2 dB
|  | 8 kHz | 0.5 dB
701. Photometry
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Iluminance meters | 70101 | 0.5 lx | 3.1 × 10 -2 | Standard Photometer
|  | (0.5 ~ 50) lx | 2.5 × 10 -2 | / KCTI-CQI-70101
|  | (50 ~ 1 500) lx | 1.4 × 10 -2
|  | (1 500 ~ 5 000) lx | 1.8 × 10 -2

## 표지(국문)-현장교정실
ISO/IEC 17025;2017 와 KS Q ISO/IEC 17025;2017 에 의한 공인인정범위
㈜한국교정기술원
충청남도 아산시 인주면 현대로 664번길 27 (인주면)
전화 : 031)497-4694 , 팩스 : 031)434-4694 , e-mail : kcti@naver.com
교 정
유효기간 만료일 : 2027. 03. 19 | [6 empty] | 인정번호 : KC07-212호
KOLAS 평가결과에 의해 다음의 교정수행에 대하여 공인됩니다.
분류번호 | 교정항목 | 현장 교정 | 분류번호 | 교정항목 | 현장 교정 | 분류번호 | 교정항목 | 현장 교정
102. 선형치수 |  |  | 104. 형상 |  |  | 10604 | 깊이 게이지,깊이마이크로미터; | Y
10201 | 볼; 강구, 표준구 등 | N | 10401 | 형상 측정기 | Y |  | 다이얼형 포함
10206 | 다이얼/실린더 게이지 시험기 | N | 10404 | 옵티컬 플랫 | N | 10609 | 지침 측미기,테스트 인디케이터 | Y
10209 | 엔드바, 마이크로미터 기준봉 | N | 10405 | 옵티컬 패러렐 | N | 10610 | 마이크로미터 헤드 | N
10210 | 길이변위계, LVDT | Y | 10406 | 평행 블록 | N | 10611 | 3점 마이크로미터 | Y
10211 | 틈새 게이지 | N | 10407 | 정밀 정반 | Y | 10612 | 내측 마이크로미터 | Y
10213 | 갭 게이지 | N | 10409 | 진원도 측정기 | Y | 10613 | 외측 마이크로미터 | Y
10214 | 게이지 블록, 비교교정 | N | 10412 | 스트레이트 엣지 | N | 10617 | 표준체 | N
10216 | 높이 게이지/측정기 | Y | 10415 | 테스트 바 | N | 203. 토크
10220 | 표준 측장기 | Y | 105. 복합형상 |  |  | 20303 | 토크 렌치 및 토크 드라이버 | Y
10223 | 전기 마이크로미터 | N | 10502 | 벤치 센터 | N | 603. 진동
10224 | 높이 마이크로미터,받침블록 | N | 10503 | 접촉식 좌표 측정기 | Y | 60301 | 진동 교정기 | N
10228 | 원통형 플러그/핀 게이지, | N | 10504 | 비접촉식 좌표 측정기 | Y | 60302 | 진동 변환기 | N
| 나사 측정용 와이어 게이지 |  | 10511 | 측정 현미경, 측정 투영기 | Y | 60303 | 진동 측정기 | N
10230 | 원통형 링 게이지 | N | 10512 | 측미 현미경 | N | 901. 화학 분석
10232 | 스텝 게이지, 캘리퍼 검사기 | N | 10514 | 테이퍼 플러그 게이지 | N | 90103 | 가스분석기 | N
10234 | 초음파식 두께 측정기 | Y | 10515 | 테이퍼 링 게이지 | N
10235 | 두께 측정용 기준 시편; 코팅형, | N | 10517 | 촉침식 표면 거칠기 측정기 | Y
| 초음파 시편 |  | 10525 | 나사 플러그 게이지 | N
10236 | 피막 두께 측정기 | Y | 10526 | 테이퍼 나사 플러그 게이지 | N
10237 | 토크 암 | N | 10527 | 나사 링 게이지 | N
103. 각도 |  |  | 10530 | 위치게이지 | N
10304 | 각도정규,콤비네이션 셋 포함 | N | 106. 기타 길이 관련량
10311 | 평형/각형/전기식 수준기 | N | 10601 | 내/외측/기어 이두께 캘리퍼, | Y
10318 | 직각도 시험기 | N |  | 캘리퍼 게이지
10320 | 정밀직각기준,직각자 | N | 10603 | 실린더/보어 게이지 | Y
주석
1. 위 기관은 고정표준실교정 및 현장교정 서비스를 제공한다.
2. 현장교정 서비스를 제공하는 기관은 "KOLAS-SR-007 현장 시험 및 교정 수행을 위한 추가기술요건"을 만족한다.
3. 인정범위 상에 "Y"가 표기된 항목은 현장교정 서비스가 가능하고, "N"이 표기된 항목은 현장교정 서비스가 불가능하다.
4. 측정불확도는 포함인자 k=2를 사용하며, 신뢰수준 약 95 %를 나타내는 확장불확도로 표현되며 정상적인 조건에서 고객에 제공될 수 있는 최소의 측정불확도를 의미한다.
5. 일반적으로, 교정성적서에 기재된 측정불확도는 교정용 표준기, 고객의 장비와 교정환경 등의 영향 때문에, 공인인정범위에 기재된 측정불확도보다 더 크게 표현됨을 유의하여야 한다.

## 교정측정능력(국문)-현장교정실
102. 선형치수
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
볼; 강구, 표준구 등 | 10201 | (5 ~ 50) mm | 0.52 + (0.0022 × l02) μm | 표준 측장기,접촉식 좌표측정기
|  |  | (l0:mm) | / KCTI-CQI-10201
다이얼/실린더 게이지 시험기 | 10206 | (0 ~ 100) mm | 0.332 + (0.0022 × l02) μm | 레이져 측정기
|  |  | (l0:mm) | 게이지블록/ KCTI-CQI-10206
엔드바, 마이크로미터 기준봉 | 10209 | (25 ~ 500) mm | 1.12 + (0.0032 × l02) μm | 전기마이크로미터
|  |  | (l0:mm) | 게이지블록/ KCTI-CQI-10209
길이변위계, LVDT | 10210 | (0 ~ 1 000) mm | 1.32 + (0.004 22 × l02) μm | 레이져 측정기
|  |  | (l0:mm) | / KCTI-CQI-10210
틈새 게이지 | 10211 | (0.01 ~ 3) mm | 2.3 μm | 벤치 마이크로미터
/ KCTI-CQI-10211
갭 게이지 | 10213 | (5 ~ 600) mm | 2.22 + (0.0062 × l02) μm | 전기마이크로미터
|  |  | (l0:mm) | / KCTI-CQI-10213
게이지 블록 , 비교교정 | 10214 | (0.5 ~ 100) mm | 752 + (1.62 × l02) nm | 게이지블록 비교기
|  |  | (l0:mm) | / KCTI-CQI-10214
높이 게이지/측정기 | 10216 | (0 ~ 1 000) mm | 0.82 + (0.0032 × l02) μm | 게이지블록,스텝게이지
|  |  | (l0:mm) | / KCTI-CQI-10216
표준 측장기 | 10220 | (0 ~ 600) mm | 0.42 + (0.002 62 × l02) μm | 레이져 측정기
|  |  | (l0:mm) | / KCTI-CQI-10220
전기 마이크로미터 | 10223 | (0 ~ 5) mm | 0.14 μm | 게이지블록
/ KCTI-CQI-10223
높이마이크로미터,받침블록 | 10224
높이 마이크로미터
블록 |  | (0 ~ 600) mm | 0.92 + (0.002 42 × l02) μm | 전기마이크로미터
|  |  | (l0:mm) | 게이지블록/ KCTI-CQI-10224
헤드 |  | (0 ~ 25) mm | 1.12 + (0.002 42 × l02) μm
(l0:mm)
원통형 플러그/핀 게이지, 나사 | 10228
측정용 와이어 게이지
원통형 플러그/핀 게이지 |  | (0.1 ~ 100) mm | 0.8 μm | 벤치 마이크로미터
표준 측장기
나사 측정용 와이어 게이지 |  | (0.1 ~ 10) mm | 0.8 μm | / KCTI-CQI-10228
원통형 링 게이지 | 10230 | (3 ~ 300) mm | 0.702 + (0.0032 × l02) μm | 표준 측장기
|  |  | (l0:mm) | / KCTI-CQI-10230
스텝 게이지, 캘리퍼 검사기 | 10232 | (0 ~ 1 000) mm | 0.52 + (0.0022 × l02) μm | 게이지블록,롱 게이지 블록
|  |  | (l0:mm) | 전기마이크로미터
/ KCTI-CQI-10232
초음파식 두께 측정기 | 10234 | (0 ~ 100) mm | 10 μm | 게이지블록,기준편
/ KCTI-CQI-10234
102. 선형치수
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
두께 측정용 기준 시편; 코팅형, | 10235
초음파 시편
코팅형 |  | (1 ~ 5 000) μm | 1.3 μm | 게이지 블록,표준 측장기
/ KCTI-CQI-10235
초음파 시편 |  | (0 ~ 200) mm | 1.02 + (0.0032 × l02) μm
(l0:mm)
피막 두께 측정기 | 10236 | (0 ~ 1.6) mm | 3.4 μm | 표준 두께 시편
/ KCTI-CQI-10236
토크암 | 10237 | (0 ~ 500) mm | 1.42 + (0.002 42 × l02) μm | 접촉식 좌표측정기
|  |  | (l0:mm) | / KCTI-CQI-10237
103. 각도
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
각도정규,콤비네이션 셋 포함 | 10304
눈금정확도 |  | (0 ~ 90)˚ | 0.9´ | 각도 눈금 원판
각도 부속품과 기준면의 각도 |  | (0 ~ 90)˚ | 1.8´ | 전기 마이크로미터
/ KCTI-CQI-10304
평형/각형/전기식 수준기 | 10311
전기식 |  | ±0.5˚ | 0.8˝ | 미소각도 설정기
평면도 |  |  | 0.7 μm | 전기 마이크로미터
직각도 |  |  | 1.9 μm | / KCTI-CQI-10311
직각도 시험기 | 10318 | (0 ~ 600) mm | 1.8 μm | 원통스퀘어/ KCTI-CQI-10318
정밀직각기준, 직각자 | 10320
직각도 |  | (0 ~ 1 000) mm | 3 μm | 원통 스퀘어,전기마이크로미터
평행도 |  | (0 ~ 1 000) mm | 1.8 μm | / KCTI-CQI-10320
104. 형상
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
형상 측정기 | 10401
세로방향 (Z축) |  | (0 ~ 100) mm | 0.37 μm | 형상 표준 시편
게이지 블록
/ KCTI-CQI-10401
가로방향 (X축) |  | (1 ~ 50) mm | 0.90 μm
각도 |  | (0 ~ 90)˚ | 4.0˚
옵티컬 플랫 | 10404
평면도 |  | (0 ~ 60) mm | 0.09 μm | 옵티컬 플랫,단색광원장치
/ KCTI-CQI-10404
옵티컬 패러렐 | 10405
평면도 |  | (12 ~ 50) mm | 0.04 μm | 게이지블록 비교기
평행도 |  | (12 ~ 50) mm | 0.07 μm | / KCTI-CQI-10405
평행 블록 | 10406
평면도 |  | (0 ~ 300) mm | 0.5 μm | 전기 마이크로미터,정밀정반
평행도 |  | (0 ~ 300) mm | 0.6 μm | 높이 게이지
짝(한쌍)의 높이 및 폭의 차이 |  | (0 ~ 1) mm | 0.7 μm | / KCTI-CQI-10406
104. 형상
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
정밀 정반 | 10407
대각선 길이 |  | (0 ~ 700) mm | 0.72 μm | 수준기
|  | (700 ~ 1 400) mm | 1.4 μm | / KCTI-CQI-10407
|  | (1 400 ~ 2 800) mm | 2.1 μm
|  | (2 800 ~ 4 000) mm | 2.6 μm
진원도 측정기 | 10409
검출기 정확도 |  | (0 ~ 300) μm | 0.22 μm | 진원도 표준 시편
스핀들 원주방향 회전 정확도 |  | (0 ~ 360)˚ | 20 nm | 진원도 표준 반구
스핀들 축방향 회전 정확도 |  | (0 ~ 360)˚ | 62 nm | / KCTI-CQI-10409
스트레이트 엣지 | 10412 | (0 ~ 500) mm | 1.5 μm | 수준기,전기 마이크로미터
/ KCTI-CQI-10412
테스트 바 | 10415 | (0 ~ 300) mm | 1.5 μm | 테스트 바
/ KCTI-CQI-10415
105. 복합형상
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
벤치 센터 | 10502
양 센터 평행도 |  | (100 ~ 300 ) mm | 3.0 μm | 게이지 블록
베드면 평면도 |  | (100 ~ 300 ) mm | 1.4 μm | 전기 마이크로미터
양 센터 높이차 |  | (100 ~ 300 ) mm | 3.0 μm | / KCTI-CQI-10502
접촉식 좌표 측정기 | 10503
X,Y,Z 지시정확도 |  | (0 ~ 1 000) mm | 0.82 + (0.0042 × l02) μm | 직각자
|  |  | (l0:mm) | 게이지 블록
대각선 |  | (0 ~ 1 000) mm | 0.82 + (0.0042 × l02) μm | 스텝 게이지
|  |  | (l0:mm) | / KCTI-CQI-10503
진직도 |  |  | 3.0 μm
직각도 |  |  | 2.0˝
비접촉식 좌표 측정기 | 10504
지시정확도 |  | (0 ~ 300) mm | 0.8 μm | 표준자,직각자
|  | (300 ~ 500) mm | 1.3 μm | / KCTI-CQI-10504
직각도(Y/X) |  | - | 0.88˝
측정 현미경, 측정 투영기 | 10511
측정 현미경 |  | (0 ~ 300) mm | 2.0 μm | 표준자
|  | (300 ~ 500) mm | 3.1 μm | / KCTI-CQI-10511
직각도 |  | (0 ~ 200) mm | 3.1 μm
측정 투영기 |  | (0 ~ 300) mm | 3.5 μm | 표준자
직각도 |  | 100 mm | 2.6 μm | / KCTI-CQI-10511
배율오차 |  | 10 배 | 0.026 %
투영면 회전 각도 |  | (0 ~ 360)˚ | 2´
측미 현미경 | 10512 | (0 ~ 10) mm | 1.1 μm | 표준자
/ KCTI-CQI-10512
105. 복합형상
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
테이퍼 플러그 게이지 | 10514
외경 |  | (0 ~ 100) mm | 1.6 μm | 게이지 블록,전기 마이크로미터
높이 |  | (0 ~ 100) mm | 2.4 μm | / KCTI-CQI-10514
반각 |  | (0 ~ 90)˚ | 0.006˚
105. 복합형상
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
테이퍼 링 게이지 | 10515
내경 |  | (0 ~ 200) mm | 1.6 μm | 표준구,전기 마이크로미터
높이 |  | (0 ~ 100) mm | 1.4 μm | / KCTI-CQI-10515
반각 |  | (0 ~ 90)˚ | 0.009˚
촉침식 표면 거칠기 측정기 | 10517
Ra |  | (0 ~ 20) μm | 0.05 μm | 표면 거칠기 시편
Rz |  | (0 ~ 50) μm | 0.16 μm | / KCTI-CQI-10517
RSm |  | (0 ~ 1) mm | 0.33 μm
단차높이 |  | (0 ~ 10) μm | 0.09 μm
나사 플러그 게이지 | 10525
유효 지름 |  | (1 ~ 100) mm | 1.7 μm | 표준측장기
피치 |  | (0 ~ 3) mm | 1.1 μm | 나사측정용 와이어 게이지
바깥 지름 |  | (0 ~ 100) mm | 1.3 μm | 비접촉식 좌표 측정기
산의 반각 |  | (0 ~ 60)˚ | 1.4´ | / KCTI-CQI-10525
테이퍼 나사 플러그 게이지 | 10526
유효지름(소,대단부) |  | (4 ~ 200) mm | 3.6 μm | 표준측장기
바깥지름(소,대단) |  | (4 ~ 200) mm | 2.2 μm | 나사측정용 와이어 게이지
피치 |  | (0.5 ~ 4) mm | 1.5 μm | 비접촉식 좌표 측정기
산의 반각 |  | (0 ~ 45)˚ | 1′ 23″ | / KCTI-CQI-10526
게이지의 길이 |  | (0 ~ 250) mm | 1.6 μm
노치의 길이 |  | (0 ~ 100) mm | 2.3 μm
나사 링 게이지 | 10527
유효 지름 |  | (4 ~ 100) mm | 1.9 μm | 표준측장기,3점 마이크로미터
안지름 |  | (4 ~ 100) mm | 2.4 μm | / KCTI-CQI-10527
위치 게이지 | 10530 | (0 ~ 300) mm | 1.12 + (0.0032 × l02) μm | 3차원 접촉식 좌표측정기
|  |  | (l0:mm) | / KCTI-CQI-10530
106. 기타 길이 및 관련량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
내/외측/기어 이두께 캘리퍼, | 10601
캘리퍼 게이지
내/외측 캘리퍼 |  | (0 ~ 1 500) mm | 102 + (0.0072 × l02) μm | 캘리퍼검사기,게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10601
캘리퍼 게이지 |  | (0 ~ 100) mm | 1.12 + (0.0072 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10601
실린더/보어 게이지 | 10603 | (0 ~ 2) mm | 0.7 μm | 다이얼게이지시험기
/ KCTI-CQI-10603
깊이 게이지,깊이마이크로미터; | 10604
다이얼형 포함
깊이 게이지,다이얼형 포함 |  | (0 ~ 300) mm | 7.82 + (0.0072 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10604
깊이 마이크로미터 |  | (0 ~ 150) mm | 1.22 + (0.0032 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10604
다이얼/디지털 게이지 | 10605 | (0 ~ 100) mm | 0.72 + (0.0122 × l02) μm | 다이얼게이지시험기
|  |  | (l0:mm) | / KCTI-CQI-10605
106. 기타 길이 및 관련량
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
지침 측미기, 테스트 인디케이터 | 10609
지침 측미기 |  | (0 ~ 1) mm | 0.7 μm | 다이얼게이지시험기
/ KCTI-CQI-10609
테스트 인디케이터 |  | (0 ~ 2) mm | 0.7 μm | 다이얼게이지시험기
/ KCTI-CQI-10609
마이크로미터 헤드 | 10610 | (0 ~ 100) mm | 1.32 + (0.0032 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10610
3점 마이크로미터 | 10611 | (3 ~ 100) mm | 1.22 + (0.0042 × l02) μm | 원통형 링 게이지
|  |  | (l0:mm) | / KCTI-CQI-10611
내측 마이크로미터 | 10612
내측 마이크로미터 |  | (5 ~ 200) mm | 1.72 + (0.0072 × l02) μm | 게이지블록
|  |  | (l0:mm) | / KCTI-CQI-10612
봉형 마이크로미터 |  | (50 ~ 500) mm | 1.22 + (0.0042 × l02) μm | 게이지블록,표준 측장기
|  |  | (l0:mm) | / KCTI-CQI-10612
외측 마이크로미터 | 10613
외측 마이크로미터 |  | (0 ~ 25) mm | 0.62 + (0.0022 × l02) μm | 게이지블록
|  | (25 ~ 125) mm | 1.02 + (0.002 42 × l02) μm | 롱 게이지블록
|  | (125 ~ 500) mm | 1.22 + (0.0052 × l02) μm | / KCTI-CQI-10613
(l0:mm)
V-앤빌 마이크로미터 |  | (2.5 ~ 25) mm | 1.3 μm | 핀 게이지
/ KCTI-CQI-10613
표준체 | 10617
선재의 지름 |  | (0.01 ~ 10) mm | 4.6 μm | 비접촉식 좌표측정기
체눈의 크기 |  | (0.01 ~ 100) mm | 2.3 μm | / KCTI-CQI-10609
203. 토크
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
토크 렌치 및 토크 드라이버 | 20303 |  |  | 토크측정기
|  | (6 ~ 60) N·m | 1.3 × 10-2 | / KCTI-CQI-20303
|  | (60 ~ 400) N·m | 1.0 × 10-2
|  | (400 ~ 1500) N·m | 1.8 × 10-2
603. 진동
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
진동 교정기 | 60301 | (20 ~ 160) Hz | 1.7 × 10-2 | 표준가속도계
|  | (160 ~ 1 250) Hz | 1.9 × 10-2 | / KCTI-CQI-60301
진동 변환기 | 60302 | 10 Hz | 3.1 × 10-2 | 표준가속도계
|  | (10 ~ 20) Hz | 2.5 × 10-2 | / KCTI-CQI-60302
|  | (20 ~ 160) Hz | 2.0 × 10-2
|  | (160 ~ 1 250) Hz | 2.2 × 10-2
|  | (1 250 ~ 2 500) Hz | 2.7 × 10-2
|  | (2 500 ~ 5 000) Hz | 3.1 × 10-2
603. 진동
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
진동 측정기 | 60303
가속도 |  | (10 ~ 160) Hz | 1.9 × 10-2 | 표준가속도계
|  | (160 ~ 2 500) Hz | 2.1 × 10-2 | / KCTI-CQI-60303
속도 |  | 10 Hz | 2.1 × 10-2
|  | (10 ~ 160) Hz | 2.0 × 10-2
|  | (160 ~ 630) Hz | 2.2 × 10-2
변위 |  | 10 Hz | 2.1 × 10-2
|  | (10 ~ 160) Hz | 2.0 × 10-2
|  | (160 ~ 315) Hz | 2.2 × 10-2
901. 화학 분석
측정량/장비 | 분류번호 | 교정범위 | 측정불확도 (신뢰수준 약 95 %) | 사용표준/측정방법 등
가스분석기 | 90103
산소(O2) |  | (0 ~ 18) cmol/mol | 2.1 × 10-2 | 표준가스
일산화탄소(CO) |  | (0 ~ 100) μmol/mol | 2.3 × 10-2 | / KCTI-CQI-90103
이산화탄소(CO2) |  | (0 ~ 5) cmol/mol | 1.7 × 10-2
메탄(CH4) |  | (0 ~ 2.5) cmol/mol | 2.1 × 10-2

## 표지(영문)현장교정실
SCOPE OF ACCREDITATION TO ISO/IEC 17025-2017 & KS Q ISO/IEC 17025-2017
Korea Calibration Technology Institute Co.,Ltd
(Inju-ri) 27, 664 Hyundai Road Inchom-myeon Asan-si Chungcheongnam-do, Republic of Korea
Phone No : 031)497-4694, Fax : 031)434-4694, e-mail : kcti@naver.com
CALIBRATION
Valid To : Mar. 19. 2027 | [7 empty] | Accreditation No. : KC07-212
In recognition of the successful completion of the KOLAS evaluation process,
accreditation is granted to this laboratory to perform the following calibrations
Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site | Field Code | Item of Calibration | on-site
102. Linear dimension |  |  | 10412 | Straight edges | N | 60303 | Vibration measuring instruments | N
10201 | Balls | N | 10415 | Test bars | N | 901. Chemical analysis
10206 | Dial/cylinder gauge testers | N | 105. Complex geometry |  |  | 90103 | Gas analyzers | N
10209 | End bars | N | 10502 | Bench centers | N
10210 | Extensometers, linear | Y | 10503 | Contact coordinate | Y
| displacement transducers |  |  | measuring machines
10211 | Filler gauges | N | 10504 | Non-contact coordinate | Y
10213 | Gap gauges | N |  | measuring machines
10214 | Gauge blocks, by comparison | N | 10511 | Measuring microscopes, | Y
10216 | Height gauges/measuring | Y |  | profile projectors
| machines |  | 10512 | Micro measuring microscopes | N
10220 | Standard measuring machines | Y | 10514 | Taper plug gauges | N
10223 | Electronic micrometers | N | 10515 | Taper ring gauges | N
10224 | Height micrometers, | N | 10517 | Stylus type roughness | Y
| Riser blocks |  |  | testers
10228 | Cylindrical plug/pin gauges | N | 10525 | Thread plug gauges | N
| thread measuring wire gauges |  | 10526 | Taper thread plug gauges | N
10230 | Cylindrical ring gauges | N | 10527 | Thread ring gauges | N
10232 | Step gauges | N | 10530 | Position gauges | N
10234 | Ultrasonic thickness gauges | Y | 106. Various dimensional
10235 | Ultrasonic/coating | N | 10601 | Inside/outside/gear tooth | Y
| thickness specimens |  |  | calipers,caliper gauges
10236 | Coating thickness testers | Y | 10603 | Cylinder/bore gauges | Y
10237 | Torque arms | N | 10604 | Depth gauges, depth | Y
103. 각도 |  |  |  | micrometers
10304 | Bevel protractors | N | 10605 | Dial/digital gauges | Y
10311 | Plate/square/electric levels | N | 10609 | Micro indicators, Test | Y
10318 | Squareness testers | N |  | indicators
10320 | Precision squares | N | 10610 | Micrometer heads | N
104. Form |  |  | 10611 | 3-point micrometers | Y
10401 | Form testers | Y | 10612 | Inside micrometers | Y
10404 | Optical flats | N | 10613 | Outside micrometers | Y
10405 | Optical parallels | N | 10617 | Standard sieves | N
10406 | Parallel blocks | N | 203. Torque
10407 | Precision surface plates | Y | 20303 | Torque wrenches/drivers | Y
10409 | Roundness measurement | Y | 603. Vibration
| instruments |  | 60301 | Vibration calibrators | N
|  |  | 60302 | Vibration transducers | N
Note
1. This laboratory provides calibration services in permanent standard laboratory and at on-site.
2. Laboratory conducts on-site calibration should meet requirements of KOLAS-SR-008.
3. On-site calibration is allowed to items with marking 'Y', not allowed to items with marking 'N'.
4. Calibration and Measurement Capability (CMC) means capabilities provided by accredited calibration laboratories. It expresses the lowest uncertainty of measurement that can be achieved during a calibration. CMC normally is quoted as an expanded uncertainty at a coverage probability of 95 %, which usually requires the use of a coverage factor of k=2.
5. Due to the calibration environment such as reference standards or customers’ facilities, it is note that uncertainty of measurement on a calibration certificate may be expressed larger than CMC on scope of accreditation in general.

## 교정측정능력(영문)현장교정실
102. Linear dimension
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Balls | 10201 | (5 ~ 50) mm | 0.52 + (0.0022 × l02) μm | Standard measuring machines
|  |  | (l0:mm) | / KCTI-CQI-10201
Dial/cylinder gauge testers | 10206 | (0 ~ 100) mm | 0.332 + (0.0022 × l02) μm | Electronic micrometer
|  |  | (l0:mm) | Gauge blocks
/ KCTI-CQI-10206
End bars | 10209 | (25 ~ 500) mm | 1.12 + (0.0032 × l02) μm | Electronic micrometer
|  |  | (l0:mm) | Gauge blocks
/ KCTI-CQI-10209
Extensometers,linear | 10210 | (0 ~ 1 000) mm | 1.32 + (0.004 22 × l02) μm | Electronic micrometer
placement transducers |  |  | (l0:mm) | Gauge blocks
/ KCTI-CQI-10210
Filler gauges | 10211 | (0.01 ~ 3) mm | 2.3 μm | Micrometers,Bench
/ KCTI-CQI-10211
Gap gauges | 10213 | (5 ~ 600) mm | 2.22 + (0.0062 × l02) μm | Electronic micrometer
|  |  | (l0:mm) | / KCTI-CQI-10213
Gauge blocks, by comparison | 10214 | (0.5 ~ 100) mm | 752 + (1.62 × l02) nm | Gauge blocks Comparator
|  |  | (l0:mm) | / KCTI-CQI-10214
Height gauges/measuring | 10216 | (0 ~ 1 000) mm | 0.82 + (0.0032 × l02) μm | Gauge blocks,Step gauge
machines |  |  | (l0:mm) | / KCTI-CQI-10216
Standard measuring machines | 10220 | (0 ~ 600) mm | 0.42 + (0.002 62 × l02) μm | Gauge blocks,
|  |  | (l0:mm) | Long Gauge blocks
/ KCTI-CQI-10220
Electronic micrometers | 10223 | (0 ~ 5) mm | 0.14 μm | Gauge blocks
/ KCTI-CQI-10223
Height micrometers,Riser blocks | 10224
Height micrometers
Block |  | (0 ~ 600) mm | 0.92 + (0.002 42 × l02) μm | Electronic micrometer
|  |  | (l0:mm) | Gauge blocks
Head |  | (0 ~ 25) mm | 1.12 + (0.002 42 × l02) μm | / KCTI-CQI-10224
(l0:mm)
Cylindrical plug/pin gauges, | 10228
thread measuring wire gauges
Cylindrical plug/pin gauges |  | (0.1 ~ 100) mm | 0.8 μm | Micrometers,Bench
Standard measuring machines
thread measuring wire gauges |  | (0.1 ~ 10) mm | 0.8 μm | / KCTI-CQI-10228
Cylindrical ring gauges | 10230 | (3 ~ 300) mm | 0.702 + (0.0032 × l02) μm | Standard measuring
|  |  | (l0:mm) | / KCTI-CQI-10230
Step gauges | 10232 | (0 ~ 1 000) mm | 0.52 + (0.0022 × l02) μm | Electronic micrometer
|  |  | (l0:mm) | Gauge blocks
/ KCTI-CQI-10232
Ultrasonic thickness gauges | 10234 | (0 ~ 100) mm | 10 μm | Gauge blocks
/ KCTI-CQI-10234
102. Linear dimension
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Ultrasonic/coating | 10235
thickness specimens
Coating |  | (1 ~ 5 000) μm | 1.3 μm | Standard measuring
machines
Ultrasonic thickness specimens |  | (0 ~ 200) mm | 1.02 + (0.0032 × l02) μm | / KCTI-CQI-10235
(l0:mm)
Coating thickness testers | 10236 | (0 ~ 1.6) mm | 3.4 μm | Standard thickness
/ KCTI-CQI-10236
Torque arms | 10237 | (0 ~ 500) mm | 1.42 + (0.00242 × l02) μm | Contact coordinate measuring
|  |  | (l0:mm) | machines
/ KCTI-CQI-10237
103. Angle
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Bevel protractors | 10304
Scale accuracy |  | (0 ~ 90)˚ | 0.9´ | Indexing tables
Angle of reference plane |  | (0 ~ 90)˚ | 1.8´ | Electronic micrometers
/ KCTI-CQI-10304
Plate/square/electric levels | 10311
Electronical |  | ±0.5˚ | 0.8˝ | Fine angle generators
Precision flat |  |  | 0.7 μm | Electronic micrometers
Right angle |  |  | 1.9 μm | / KCTI-CQI-10311
Squareness testers | 10318 | (0 ~ 600) mm | 1.8 μm | Cylindrical squares
/ KCTI-CQI-10318
Precision squares | 10320 |  |  | Cylindrical squares
Right angle |  | (0 ~ 1 000) mm | 3 μm | Electronic micrometers
Parallelism |  | (0 ~ 1 000) mm | 1.8 μm | / KCTI-CQI-10320
104. Form
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Form testers | 10401
Vertical (Z) |  | (0 ~ 100) mm | 0.37 μm | Form standard thickness
Gauge blocks
/ KCTI-CQI-10401
Horizontal (X) |  | (1 ~ 50) mm | 0.90 μm
Angle |  | (0 ~ 90)˚ | 4.0˚
Optical flats | 10404
Floor plan |  | (0 ~ 60) mm | 0.09 μm | Optical flats
/ KCTI-CQI-10404
Optical parallels | 10405
Floor plan |  | (12 ~ 50) mm | 0.04 μm | Gauge blocks Comparator
Parallelism |  | (12 ~ 50) mm | 0.07 μm | / KCTI-CQI-10405
Parallel blocks | 10406
Floor plan |  | (0 ~ 300) mm | 0.5 μm | Electronic micrometer
Parallelism |  | (0 ~ 300) mm | 0.6 μm | Gauge blocks
Differences in height and |  | (0 ~ 1) mm | 0.7 μm | / KCTI-CQI-10406
width of pairs
104. Form
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Precision surface plates | 10407
Diagonal length |  | (0 ~ 700) mm | 0.72 μm | Electrical levels
|  | (700 ~ 1 400) mm | 1.4 μm | / KCTI-CQI-10407
|  | (1 400 ~ 2 800) mm | 2.1 μm
|  | (2 800 ~ 4 000) mm | 2.6 μm
Roundness measurement | 10409
instruments |  |  |  | Roundness standard
Detector accuracy |  | (0 ~ 300) μm | 0.22 μm | thickness
Spindle Rotation accuracy |  | (0 ~ 360)˚ | 20 nm | Roundness standard radius
Spindle shaft accuracy |  | (0 ~ 360)˚ | 62 nm | / KCTI-CQI-10409
Straight edges | 10412 | (0 ~ 500) mm | 1.5 μm | Electrical levels
/ KCTI-CQI-10412
Test bar | 10415 | (0 ~ 300) mm | 1.5 μm | Test bar
/ KCTI-CQI-10415
105. Complex geometry
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Bench centers | 10502
Amount of parallelism Center |  | (100 ~ 300 ) mm | 3.0 μm | Electronic micrometer
Bed surface plan |  | (100 ~ 300 ) mm | 1.4 μm | Test Bar
Sheep Center height difference |  | (100 ~ 300 ) mm | 3.0 μm | / KCTI-CQI-10502
Contact coordinate measuring | 10503
machines
X,Y,Z accuracy |  | (0 ~ 1 000) mm | 0.82 + (0.0042 × l02) μm | Right angle rule
|  |  | (l0:mm) | Gauge blocks
Accuray |  | (0 ~ 1 000) mm | 0.82 + (0.0042 × l02) μm | Step gauge
|  |  | (l0:mm) | / KCTI-CQI-10503
Straightness |  |  | 3.0 μm
Right angle |  |  | 2.0˝
Non-contact coordinate | 10504
measuring machines
Accuracy instructions |  | (0 ~ 300) mm | 0.8 μm | Standard scales
|  | (300 ~ 500) mm | 1.3 μm | Right angle rule
Squareness(Y/X) |  | - | 0.88˝ | / KCTI-CQI-10504
Measuring microscopes, | 10511
Profile projectors
Measuring microscopes |  | (0 ~ 300) mm | 2.0 μm | Standard scales
|  | (300 ~ 500) mm | 3.1 μm | / KCTI-CQI-10511
Right angle |  | (0 ~ 200) mm | 3.1 μm
Profile projectors |  | (0 ~ 300) mm | 3.5 μm | Standard scales
Right angle |  | 100 mm | 2.6 μm
Magnification error |  | 10 magnification | 0.026 %
Projection plant rotation angle |  | (0 ~ 360)˚ | 2´ | / KCTI-CQI-10511
Micro measuring microscopes | 10512 | (0 ~ 10) mm | 1.1 μm | Standard scales
/ KCTI-CQI-10512
105. Complex geometry
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Taper plug gauges | 10514
external diameter |  | (0 ~ 100) mm | 1.6 μm | Electronic micrometer
height |  | (0 ~ 100) mm | 2.4 μm | / KCTI-CQI-10514
half angle |  | (0 ~ 90)˚ | 0.006˚
Taper ring gauges | 10515
internal diameter |  | (0 ~ 200) mm | 1.6 μm | Electronic micrometer
height |  | (0 ~ 100) mm | 1.4 μm | Master ball
half angle |  | (0 ~ 90)˚ | 0.009˚ | / KCTI-CQI-10515
Stylus type roughness testers | 10517
Ra |  | (0 ~ 20) μm | 0.05 μm | Stylus Type Roughness
Rz |  | (0 ~ 50) μm | 0.16 μm | / KCTI-CQI-10517
RSm |  | (0 ~ 1) mm | 0.33 μm
Step height |  | (0 ~ 10) μm | 0.09 μm
Thread plug gauges | 10525
Major diameter |  | (1 ~ 100) mm | 1.7 μm | Micrometers,Bench
Pitch diameter |  | (0 ~ 3) mm | 1.1 μm | wire gauges
Pitch |  | (0 ~ 100) mm | 1.3 μm | / KCTI-CQI-10525
Thread half angle |  | (0 ~ 60)˚ | 1.4´
Taper thread plug gauges | 10526
Smallend pitch diameter |  | (4 ~ 200) mm | 3.6 μm | Gauge blocks,Standard
Bigend pitch diameter |  | (4 ~ 200) mm | 2.2 μm | measuring machines
Pitch |  | (0.5 ~ 4) mm | 1.5 μm | Electronic micrometer
Half angle of thread |  | (0 ~ 45)˚ | 1′ 23″ | Bench wire gauges
Gauge length |  | (0 ~ 250) mm | 1.6 μm | highe master,Measuring
Notch length |  | (0 ~ 100) mm | 2.3 μm | microscopes
/ KCTI-CQI-10526
Thread ring gauges | 10527 |  |  | Standard measuring
Pitch diameter |  | (4 ~ 100) mm | 1.9 μm | machines,3-point
inside diameter |  | (4 ~ 100) mm | 2.4 μm | Micrometers
/ KCTI-CQI-10527
Position gauges | 10530 | (0 ~ 300) mm | 1.12 + (0.0032 × l02) μm | Contact coordinate measuring
|  |  | (l0:mm) | machines
/ KCTI-CQI-10530
106. Various dimensional
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Inside/outside/gear tooth | 10601
calipers,caliper gauges
Inside/Outside calipers |  | (0 ~ 1 500) mm | 102 + (0.0072 × l02) μm | Step gauges
|  |  | (l0:mm) | / KCTI-CQI-10601
Caliper gauges |  | (0 ~ 100) mm | 1.12 + (0.0072 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10601
Cylinder/Bore gauges | 10603 | (0 ~ 2) mm | 0.7 μm | Dial gauge testers
/ KCTI-CQI-10603
106. Various dimensional
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Depth gauges,Depth micrometers | 10604
Depth gauges |  | (0 ~ 300) mm | 7.82 + (0.0072 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10604
Depth micrometers |  | (0 ~ 150) mm | 1.22 + (0.0072 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10604
Dial/Digital gauges | 10605 | (0 ~ 100) mm | 0.72 + (0.0122 × l02) μm | Dial gauge testers
|  |  | (l0:mm) | / KCTI-CQI-10605
Micro indicators, Test | 10609
indicators
Micro indicators |  | (0 ~ 1) mm | 0.7 μm | Dial gauge testers
/ KCTI-CQI-10609
Test indicators |  | (0 ~ 2) mm | 0.7 μm | Dial gauge testers
/ KCTI-CQI-10609
Micrometer heads | 10610 | (0 ~ 100) mm | 1.32 + (0.0032 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10610
3-point micrometers | 10611 | (3 ~ 100) mm | 1.22 + (0.0042 × l02) μm | Ring Gauge
|  |  | (l0:mm) | / KCTI-CQI-10611
Inside micrometers | 10612
Inside micrometers |  | (5 ~ 200) mm | 1.72 + (0.0072 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10612
Micrometers, bar type |  | (50 ~ 500) mm | 1.22 + (0.0042 × l02) μm | Gauge blocks
|  |  | (l0:mm) | / KCTI-CQI-10612
Outside micrometers | 10613
Outside micrometers |  | (0 ~ 25) mm | 0.62 + (0.0022 × l02) μm | Gauge blocks
(l0:mm)
|  | (25 ~ 125) mm | 1.02 + (0.002 42 × l02) μm | / KCTI-CQI-10613
(l0:mm)
|  | (125 ~ 500) mm | 1.22 + (0.0052 × l02) μm
(l0:mm)
Micrometers, V-anvil |  | (2.5 ~ 25) mm | 1.3 μm | Pin Gauge
/ KCTI-CQI-10613
Standard sieves | 10617
Diameter of wire rod |  | (0.01 ~ 10) mm | 4.6 μm | Measuring microscopes
sieve scale |  | (0.01 ~ 100) mm | 2.3 μm | / KCTI-CQI-10617
203. Torque
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Torque wrenches/drivers | 20303 |  |  | Torque calibration machine
|  | (6 ~ 60) N·m | 1.3 × 10-2 | / KCTI-CQI-20303
|  | (60 ~ 400) N·m | 1.0 × 10-2
|  | (400 ~ 1500) N·m | 1.8 × 10-2
603. Vibration
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Vibration calibrators | 60301 | (20 ~ 160) Hz | 1.7 × 10-2 | Standard accelerometer
|  | (160 ~ 1 250) Hz | 1.9 × 10-2 | / KCTI-CQI-60301
Vibration transducers | 60302 | 10 Hz | 3.1 × 10-2 | Standard accelerometer
|  | (10 ~ 20) Hz | 2.5 × 10-2 | / KCTI-CQI-60302
|  | (20 ~ 160) Hz | 2.0 × 10-2
|  | (160 ~ 1 250) Hz | 2.2 × 10-2
|  | (1 250 ~ 2 500) Hz | 2.7 × 10-2
|  | (2 500 ~ 5 000) Hz | 3.1 × 10-2
Vibration measuring | 60303
instruments
Acceleration |  | (10 ~ 160) Hz | 1.9 × 10-2 | Standard accelerometer
|  | (160 ~ 2 500) Hz | 2.1 × 10-2 | / KCTI-CQI-60303
Velocity |  | 10 Hz | 2.1 × 10-2
|  | (10 ~ 160) Hz | 2.0 × 10-2
|  | (160 ~ 630) Hz | 2.2 × 10-2
Displacement |  | 10 Hz | 2.1 × 10-2
|  | (10 ~ 160) Hz | 2.0 × 10-2
|  | (160 ~ 315) Hz | 2.2 × 10-2
901. Chemical analysis
Measured Quantity Instrument or Gauge | Field code | Range | Measurement uncertainty (The Confidence Level is about 95 %) | Standard/Method of Measurement etc.
Gas analyzers | 90103
oxygen |  | (0 ~ 18) cmol/mol | 2.1 × 10-2 | Standard Gas
Carbon monoxide(CO) |  | (0 ~ 100) μmol/mol | 2.3 × 10-2 | / KCTI-CQI-90103
Carbon dioxide(CO₂) |  | (0 ~ 5) cmol/mol | 1.7 × 10-2
Methane(CH₄) |  | (0 ~ 2.5) cmol/mol | 2.1 × 10-2